## SM-RI-X Tubine 20 bar
Institute
person in charge
Date
Environmental conditions
Mean lab. Temp. |  |  | °C
Mean Baro. |  | 35 | bar
manufacture | model | serial number | size | G-Range | Qmax [m3/h] | Qmin [m3/h] | Gas composition
Honeywell | SM-RI-X 4" | 3.130.200S | DN 100 ^ 4" | 250 | 400 | 20 | CO2
Testgas | Date | Nominel flowrate | Average temperature | average pressure | Real flow at MUT | Indicated flow MUT | deviation | Re | Average indicated flow | Average deviaition | standard deviation of the mean
|  | [m3/h] | [°C] | [bar] | [m3/h] | [m3/h] | [%] | [-] | [m3/h] | [%] | [%]
CO2 |  | 20
CO2 |  | 20
CO2 |  | 20
CO2 |  | 40
CO2 |  | 40
CO2 |  | 40
CO2 |  | 80
CO2 |  | 80
CO2 |  | 80
CO2 |  | 160
CO2 |  | 160
CO2 |  | 160
CO2 |  | 280
CO2 |  | 280
CO2 |  | 280
CO2 |  | 400
CO2 |  | 400
CO2 |  | 400

## SM-RI-X Tubine 30 bar
Institute
person in charge
Date
Environmental conditions
Mean lab. Temp. |  |  | °C
Mean Baro. |  | 1 | atm
manufacture | model | serial number | size | G-Range | Qmax [m3/h] | Qmin [m3/h] | Gas composition
Honeywell | SM-RI-X 4" | 3.130.200S | DN 100 ^ 4" | 250 | 400 | 20 | Air
Testgas | Date | Nominel flowrate | Average temperature | average pressure | Real flow at MUT | Indicated flow MUT | deviation | Re | Average indicated flow | Average deviaition | standard deviation of the mean
|  | [m3/h] | [°C] | [bar] | [m3/h] | [m3/h] | [%] | [-] | [m3/h] | [%] | [%]
Air |  | 20
Air |  | 20
Air |  | 20
Air |  | 40
Air |  | 40
Air |  | 40
Air |  | 80
Air |  | 80
Air |  | 80
Air |  | 160
Air |  | 160
Air |  | 160
Air |  | 280
Air |  | 280
Air |  | 280
Air |  | 400
Air |  | 400
Air |  | 400

## Emerson 3” Coriolis 30 bar.
Institute |  | FORCE Technology
person in charge |  | Kurt Rasmussen
Date |  | 2025-06-03
Environmental conditions
Mean lab. Temp. |  | 21.7 | °C
Mean Baro. |  | 999.8 | mbar
manufacture | model | serial number | size | Qmax [m3/h] | Qmin [m3/h] |  | Gas composition
Emerson | CMF300M | 21046486 | DN80 ^ 3" | 400 | 20 |  | NG
Testgas | Date | Nominel flowrate | Average temperature | average pressure | Real flow at MUT |  | Indicated flow MUT |  | deviation | Reynolds number | Average indicated flow | Average deviaition | standard deviation of the mean
|  | [m3/h] | [°C] | [bar a] | [m3/h] | [kg/h] | [m3/h] | [kg/h] | [%] | [] | [kg/h] | [%] | [%]
NG | 2025-06-03 | 20 | 21.3527 | 31.0133 |  | 439.2705 |  | 435.3079 | -0.90209
NG | 2025-06-03 | 20 | 21.3651 | 31.0029 |  | 439.1649 |  | 435.3118 | -0.87735
NG | 2025-06-03 | 20 | 21.3819 | 30.9956 |  | 438.9837 |  | 435.2763 | -0.84453 |  | 439.1397 | -0.87466 | 0.01667
NG | 2025-06-03 | 40 | 21.3169 | 31.0132 |  | 878.7293 |  | 873.7667 | -0.56475
NG | 2025-06-03 | 40 | 21.3191 | 31.0082 |  | 878.5858 |  | 874.7667 | -0.4347
NG | 2025-06-03 | 40 | 21.3334 | 31.0019 |  | 878.3334 |  | 874.0333 | -0.48958 |  | 878.5495 | -0.49634 | 0.037695
NG | 2025-06-03 | 80 | 21.3109 | 31.0225 |  | 1977.9224 |  | 1975.7318 | -0.11075
NG | 2025-06-03 | 80 | 21.2989 | 31.0183 |  | 1757.9329 |  | 1754.6117 | -0.18892
NG | 2025-06-03 | 80 | 21.2998 | 31.0163 |  | 1757.7979 |  | 1755.3088 | -0.1416 |  | 1831.2177 | -0.14709 | 0.022733
NG | 2025-06-03 | 160 | 22.2366 | 31.1059 |  | 3512.9243 |  | 3507.3214 | -0.15949
NG | 2025-06-03 | 160 | 22.222 | 31.102 |  | 3512.6659 |  | 3506.9678 | -0.16221
NG | 2025-06-03 | 160 | 22.2382 | 31.1003 |  | 3512.2984 |  | 3507.9536 | -0.1237 |  | 3512.6295 | -0.14847 | 0.012409
NG | 2025-06-03 | 280 | 22.2554 | 31.109 |  | 6147.9851 |  | 6140.8418 | -0.11619
NG | 2025-06-03 | 280 | 22.2691 | 31.1139 |  | 6148.5271 |  | 6146.3455 | -0.035482
NG | 2025-06-03 | 280 | 22.2961 | 31.1199 |  | 6149.311 |  | 6143.5636 | -0.093464 |  | 6148.6077 | -0.081711 | 0.024028
NG | 2025-06-03 | 400 | 22.01 | 31.097 |  | 8787.5829 |  | 8788.575 | 0.01129 |  | 7028.4737 | -0.039219
[11 empty] | 439.1397 | -0.87466
[11 empty] | 878.5495 | -0.49634
[11 empty] | 1831.2177 | -0.14709
[11 empty] | 3512.6295 | -0.14847
[11 empty] | 6148.6077 | -0.081711
[11 empty] | 7028.4737 | -0.039219
| 2025-06-03 | 20 | 21.3527 | 31.0133 |  | 439.2705 |  | 435.3079 | -0.90209
| 2025-06-03 | 20 | 21.3651 | 31.0029 |  | 439.1649 |  | 435.3118 | -0.87735
| 2025-06-03 | 20 | 21.3819 | 30.9956 |  | 438.9837 |  | 435.2763 | -0.84453
| 2025-06-03 | 40 | 21.3169 | 31.0132 |  | 878.7293 |  | 873.7667 | -0.56475
| 2025-06-03 | 40 | 21.3191 | 31.0082 |  | 878.5858 |  | 874.7667 | -0.4347
| 2025-06-03 | 40 | 21.3334 | 31.0019 |  | 878.3334 |  | 874.0333 | -0.48958
| 2025-06-03 | 80 | 21.3109 | 31.0225 |  | 1977.9224 |  | 1975.7318 | -0.11075
| 2025-06-03 | 80 | 21.2989 | 31.0183 |  | 1757.9329 |  | 1754.6117 | -0.18892
| 2025-06-03 | 80 | 21.2998 | 31.0163 |  | 1757.7979 |  | 1755.3088 | -0.1416
| 2025-06-03 | 160 | 22.2366 | 31.1059 |  | 3512.9243 |  | 3507.3214 | -0.15949
| 2025-06-03 | 160 | 22.222 | 31.102 |  | 3512.6659 |  | 3506.9678 | -0.16221
| 2025-06-03 | 160 | 22.2382 | 31.1003 |  | 3512.2984 |  | 3507.9536 | -0.1237
| 2025-06-03 | 280 | 22.2554 | 31.109 |  | 6147.9851 |  | 6140.8418 | -0.11619
| 2025-06-03 | 280 | 22.2691 | 31.1139 |  | 6148.5271 |  | 6146.3455 | -0.035482
| 2025-06-03 | 280 | 22.2961 | 31.1199 |  | 6149.311 |  | 6143.5636 | -0.093464
| 2025-06-03 | 400 | 22.01 | 31.097 |  | 8787.5829 |  | 8788.575 | 0.01129

## Emerson 3” Coriolis 20 bar
Institute |  | FORCE Technology
person in charge |  | Kurt Rasmussen
Date |  | 2025-06-03
Environmental conditions
Mean lab. Temp. |  | 21.64 | °C
Mean Baro. |  | 1007 | mbar
manufacture | model | serial number | size | Qmax [m3/h] | Qmin [m3/h] |  | Gas composition
Emerson | CMF300M | 21046486 | DN80 ^ 3" | 400 | 20 |  | NG
Testgas | Date | Nominel flowrate | Average temperature | average pressure | Real flow at MUT |  | Indicated flow MUT |  | deviation | Reynolds number | Average indicated flow | Average deviaition | standard deviation of the mean
|  | [m3/h] | [°C] | [bar a] | [m3/h] | [kg/h] | [m3/h] | [kg/h] | [%] | [] | [kg/h] | [%] | [%]
NG | 2025-06-03 | 20 | 21.4007 | 21.0233 |  | 292.3832 |  | 288.9411 | -1.1773
NG | 2025-06-03 | 20 | 21.4255 | 21.0301 |  | 292.4626 |  | 289.0547 | -1.1652 |  | 292.4229 | -1.1712 | 0.0049079
NG | 2025-06-03 | 40 | 21.3992 | 21.0265 |  | 584.8912 |  | 582.6067 | -0.3906
NG | 2025-06-03 | 40 | 21.3509 | 21.0235 |  | 584.9368 |  | 583.3333 | -0.27413 |  | 584.914 | -0.33237 | 0.27965
NG | 2025-06-03 | 80 | 21.2711 | 21.0329 |  | 1170.7551 |  | 1170.3318 | -0.036157
NG | 2025-06-03 | 80 | 21.232 | 21.0398 |  | 1171.3848 |  | 1170.9521 | -0.036934 |  | 1171.0699 | -0.036546 | 0.079197
NG | 2025-06-03 | 160 | 21.169 | 21.0452 |  | 2343.9544 |  | 2346.2 | 0.095804
NG | 2025-06-03 | 160 | 21.1818 | 21.0492 |  | 2344.3317 |  | 2344.8128 | 0.020524 |  | 2344.143 | 0.058164 | 0.038433
NG | 2025-06-03 | 280 | 21.1549 | 21.0513 |  | 4103.3446 |  | 4106.8606 | 0.085686
NG | 2025-06-03 | 280 | 21.1854 | 21.0566 |  | 4104.0007 |  | 4105.0667 | 0.025973 |  | 4103.6727 | 0.05583 | 0.020872
NG | 2025-06-03 | 400 | 20.9987 | 21.0487 |  | 5865.0293 |  | 5880.9333 | 0.27117 |  | 5865.0293 | 0.27117 | 0.07382
[11 empty] | 292.4229 | -1.1712
[11 empty] | 584.914 | -0.33237
[11 empty] | 1171.0699 | -0.036546
[11 empty] | 2344.143 | 0.058164
[11 empty] | 4103.6727 | 0.05583
CO2 30bar
[11 empty] | 2747.4132 | -0.23445
[11 empty] | 5490.9417 | -0.26293
[11 empty] | 10981.2118 | -0.23543
[11 empty] | 19167.4306 | 0.033136
CO2 20 bar
[11 empty] | 860.8179 | -0.28236
[11 empty] | 1721.5929 | -0.29291
[11 empty] | 3443.0653 | -0.33436
[11 empty] | 6875.7713 | -0.19808
[11 empty] | 11999.7219 | -0.054042