## Data
[91 empty] | (-) | (%)
| VSL | [23 empty] | INRIM | [21 empty] | NEL | [42 empty] | Total test points | 51
[90 empty] | Above 1 | 16 | 31.3725
[90 empty] | Above 1.2 | 13 | 25.4902
[25 empty] | Notes: reported data refer to Totalized Volume; ambient temperature ranged from 292.8 K to 293 K; Ambient pressure ranged from 97930 Pa to 99350 Pa (measurements taken in different days) | [64 empty] | Between 1 and 1.2 | 3 | 5.8824
[90 empty] | Below equal 1 | 35 | 68.6275
[21 empty] | 0.21467 | [19 empty] | 0.60016 | [25 empty] | 0.49178
| PGM meter | [23 empty] | PGM meter | [21 empty] | PGM meter | [23 empty] | VSL |  | INRIM |  | NEL | [5 empty] | VSL |  |  | INRIM |  |  | NEL
Date and time | Fluid | TestPoint Timestamp | Test Point | Reference time | Reference volume | Q_Reference (corrected to MUT conditions) | p_abs_MUT | T_MUT | rho_MUT | rho_MUT | Q_MUT | MUT time | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average | Fluid | TestPoint Timestamp | Test Point | Test time | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | Q_MUT | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average |  | Fluid | TestPoint Timestamp | Nozzle | Test Point | Test time | Q_Reference | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | Q_MUT | Q_MUT | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | Q ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
| - | - | - | s | ml | m3/h | Pa | °C | kg/m³ | kg/m3 | m3/h | s | % | % | % |  |  |  | % | % | % | m3 | m3/h | m3/h | - | - | - | s | m3/h | Pa | °C | kg/m³ | m3/h | % | % | % |  |  |  | % | % | % | m3 | m3/h | m3/h |  | - | - | - | - | s | m3/h | m3 | Pa | °C | kg/m³ | m3/h | m3 | % | % | % |  |  |  | % | % | % | m3 | m3/h | m3/h | [6 empty] | m3/h | % |  |  |  |  | % |  |  | % |  |  | %
2024-07-03 15:20 | N2 |  |  | 73.4785 | 2064.1923 | 0.101 | 101424.4003 | 20.3378 |  | 1.1646 | 0.10046 | 73.5618 | -0.53099 | 0.2 | -0.57255 | 0.045198 | 3 | 4.3027 | 0.11228 | 0.2 | 0.22936 | 0.10101 | 0.101 | 0.10101 | N2 |  |  | 1907.3802 | 0.094665 | 98851.873 | 19.815 | 1.1369 | 0.095061 | 0.41835 | 0.1 | -0.15958 | 0.51168 | 3 | 4.3027 | 1.2711 | 0.1 | 1.275 | 0.097142 |  | 0.097142 |  | N2 | 10/03/2025 15:29:11 | RS8 | 54 | 1211.2336 | 0.12443 | 0.041866 | 1112933.9726 | 19.1964 | 12.8582 | 0.12408 | 0.041748 | -0.28074 | 0.39 | -0.20334 | 0.26145 | 3 | 4.3027 | 0.64947 | 0.4 | 0.76276 | 0.041955 |  | 0.12468 | -10.8837 | 19.0092 | -0.098165 | 0.61513 | -0.3495 | 1.7188 | 0.10761 | -0.53091 | 21.3431 | 0.21646 | -0.041635 | 0.075846 | 0.54894 | 0.37133 | 1.2565 | 0.29552 | 0.32757 | 0.7314 | 0.44787
2024-07-03 15:22 | N2 |  |  | 73.48 | 2064.1923 | 0.101 | 101419.5354 | 20.3361 |  | 1.1646 | 0.10037 | 73.6302 | -0.62067 | 0.2 | [8 empty] | 0.101 |  | N2 |  |  | 1878.018 | 0.097739 | 99006.9441 | 19.9615 | 1.1381 | 0.097404 | -0.34217 | 0.1 | [11 empty] | N2 | 10/03/2025 15:49:46 | RS8 | 55 | 1211.7684 | 0.12476 | 0.041996 | 1143001.2217 | 18.6172 | 13.2334 | 0.12487 | 0.042033 | 0.088065 | 0.39
2024-07-03 15:23 | N2 |  |  | 73.4545 | 2064.1923 | 0.10103 | 101419.3448 | 20.3355 |  | 1.1646 | 0.10046 | 73.5642 | -0.56599 | 0.2 | [8 empty] | 0.10103 |  | N2 |  |  | 1894.7964 | 0.099022 | 98848.2994 | 19.989 | 1.1361 | 0.098473 | -0.55493 | 0.1 | [11 empty] | N2 | 10/03/2025 16:11:13 | RS8 | 56 | 1211.1956 | 0.12485 | 0.042005 | 1150746.7211 | 18.126 | 13.3465 | 0.12433 | 0.041829 | -0.41735 | 0.39
2024-07-03 15:02 | N2 |  |  | 41.5575 | 3474.9427 | 0.30045 | 101503.52 | 20.3429 |  | 1.1655 | 0.30081 | 41.5944 | 0.12294 | 0.2 | 0.13849 | 0.016309 | 3 | 4.3027 | 0.040514 | 0.2 | 0.20406 | 0.30041 | 0.30045 | 0.30041 | N2 |  |  | 623.8914 | 0.30503 | 98842.348 | 19.8285 | 1.1367 | 0.30493 | -0.031184 | 0.1 | 0.035265 | 0.41735 | 3 | 4.3027 | 1.0368 | 0.1 | 1.0416 | 0.30724 |  | 0.30724 |  | N2 | 10/03/2025 13:14:22 | RS8 | 48 | 1210.91 | 0.28805 | 0.09689 | 1090740.5483 | 18.0221 | 12.6536 | 0.28938 | 0.097337 | 0.46132 | 0.39 | 0.4269 | 0.035882 | 3 | 4.3027 | 0.089135 | 0.4 | 0.40981 | 0.093657 |  | 0.2883 | 3.3257 | 24.0146 | 0.032506 | 0.92178 | 2.5419 | 5.9543 | 0.29865 | 0.191 | 30.8907 | 0.17992 | -0.052513 | 0.096276 | 0.54544 | -0.15574 | 1.0259 | 0.1518 | 0.2359 | 0.3682 | 0.64068
2024-07-03 15:03 | N2 |  |  | 41.5635 | 3474.9427 | 0.30039 | 101498.4343 | 20.3421 |  | 1.1655 | 0.30086 | 41.5884 | 0.15546 | 0.2 | [8 empty] | 0.30039 |  | N2 |  |  | 595.581 | 0.3082 | 98981.8249 | 19.952 | 1.1378 | 0.30714 | -0.34487 | 0.1 | [11 empty] | N2 | 10/03/2025 13:35:20 | RS8 | 49 | 1210.9026 | 0.28833 | 0.096982 | 1090383.901 | 18.129 | 12.6447 | 0.28956 | 0.097398 | 0.42967 | 0.39
2024-07-03 15:04 | N2 |  |  | 41.5615 | 3474.9427 | 0.30041 | 101501.4735 | 20.3422 |  | 1.1655 | 0.30082 | 41.3508 | 0.13706 | 0.2 | [8 empty] | 0.30041 |  | N2 |  |  | 590.3382 | 0.30851 | 98864.7611 | 19.9755 | 1.1364 | 0.30999 | 0.48185 | 0.1 | [11 empty] | N2 | 10/03/2025 13:56:11 | RS8 | 50 | 1086.7304 | 0.28853 | 0.0871 | 1089587.6707 | 18.2307 | 12.6309 | 0.28966 | 0.087439 | 0.38971 | 0.39
2024-07-03 14:45 | N2 |  |  | 63.3345 | 10615.1509 | 0.60144 | 101719.9324 | 20.3485 |  | 1.168 | 0.60321 | 63.4412 | 0.29357 | 0.2 | 0.30821 | 0.013128 | 3 | 4.3027 | 0.032612 | 0.2 | 0.20264 | 0.60132 | 0.60144 | 0.60132 | N2 |  |  | 444.5898 | 0.60006 | 98844.285 | 19.827 | 1.1367 | 0.59907 | -0.16527 | 0.1 | 0.017677 | 0.162 | 3 | 4.3027 | 0.40243 | 0.1 | 0.41467 | 0.60124 |  | 0.60124 |  | N2 | 10/03/2025 14:24:38 | RS8 | 51 | 1211.0701 | 0.58066 | 0.19534 | 1081410.0909 | 19.9502 | 12.4602 | 0.58545 | 0.19695 | 0.82491 | 0.39 | 0.72781 | 0.11832 | 3 | 4.3027 | 0.29393 | 0.4 | 0.49638 | 0.1958 |  | 0.58203 | 7.5058 | 24.3525 | 0.1028 | 5.8155 | 2.9539 | 4.0586 | 0.59486 | 0.3086 | 34.2266 | 0.17093 | -0.00038986 | 0.10884 | 0.0035819 | -0.29093 | 0.3778 | 0.77005 | 0.41921 | 0.46602 | 0.89955
2024-07-03 14:46 | N2 |  |  | 63.3505 | 10615.1509 | 0.60127 | 101712.6277 | 20.3469 |  | 1.1679 | 0.60319 | 63.3222 | 0.31894 | 0.2 | [8 empty] | 0.60127 |  | N2 |  |  | 303.0354 | 0.6011 | 98966.6175 | 19.9295 | 1.1377 | 0.60196 | 0.14294 | 0.1 | [11 empty] | N2 | 10/03/2025 14:45:14 | RS8 | 52 | 1211.4002 | 0.5823 | 0.19594 | 1079347.6177 | 20.5948 | 12.4083 | 0.58674 | 0.19744 | 0.76251 | 0.39
2024-07-03 14:48 | N2 |  |  | 63.3565 | 10615.1509 | 0.60123 | 101722.4998 | 20.3467 |  | 1.168 | 0.60311 | 63.4516 | 0.31213 | 0.2 | [8 empty] | 0.60123 |  | N2 |  |  | 304.0842 | 0.60256 | 98878.8768 | 19.9685 | 1.1366 | 0.60301 | 0.07536 | 0.1 | [11 empty] | N2 | 10/03/2025 15:05:54 | RS8 | 53 | 1210.8234 | 0.58312 | 0.19613 | 1077740.0201 | 20.8083 | 12.3806 | 0.5866 | 0.1973 | 0.59603 | 0.39
2024-07-03 14:33 | N2 |  |  | 46.2285 | 13030.4958 | 1.0092 | 101967.0581 | 20.3589 |  | 1.1708 | 1.0128 | 46.3088 | 0.36105 | 0.2 | 0.35578 | 0.0060871 | 3 | 4.3027 | 0.015121 | 0.2 | 0.20057 | 1.0093 | 1.0092 | 1.0093 | N2 |  |  | 191.8878 | 1.0035 | 98883.4253 | 19.964 | 1.1366 | 1.0078 | 0.4267 | 0.1 | 0.21188 | 0.19052 | 3 | 4.3027 | 0.47327 | 0.1 | 0.48372 | 1.0045 |  | 1.0045 |  | N2 | 10/03/2025 07:24:01 | RS8 | 36 | 801.0439 | 0.91347 | 0.20326 | 1141952.0072 | 20.0041 | 13.1568 | 0.91847 | 0.20437 | 0.54655 | 0.39 | 0.37485 | 0.1586 | 3 | 4.3027 | 0.39398 | 0.4 | 0.56145 | 0.20348 |  | 0.91338 | 8.844 | 24.8579 | 0.9055 | 4.2737 | 1.1891 | 3.1724 | 0.97573 | 0.33861 | 32.304 | 0.17594 | 0.017166 | 0.096295 | 0.17826 | -0.12674 | 0.45059 | 0.28127 | 0.036232 | 0.53317 | 0.067957
2024-07-03 14:34 | N2 |  |  | 46.2225 | 13030.4958 | 1.0093 | 101964.9877 | 20.3571 |  | 1.1708 | 1.0129 | 46.2326 | 0.35717 | 0.2 | [8 empty] | 1.0093 |  | N2 |  |  | 186.645 | 1.0037 | 98846.0245 | 19.824 | 1.1368 | 1.0044 | 0.063431 | 0.1 | [11 empty] | N2 | 10/03/2025 07:38:42 | RS8 | 37 | 802.1168 | 0.91346 | 0.20353 | 1141189.4584 | 20.1275 | 13.1423 | 0.9166 | 0.20423 | 0.34417 | 0.39
2024-07-03 14:36 | N2 |  |  | 46.2165 | 13030.4958 | 1.0095 | 101960.0071 | 20.3566 |  | 1.1707 | 1.013 | 46.0828 | 0.34912 | 0.2 | [8 empty] | 1.0095 |  | N2 |  |  | 181.4022 | 1.0061 | 98965.4422 | 19.9245 | 1.1377 | 1.0076 | 0.14549 | 0.1 | [11 empty] | N2 | 10/03/2025 07:52:46 | RS8 | 38 | 802.8189 | 0.91321 | 0.20365 | 1140759.5631 | 20.1397 | 13.1368 | 0.91534 | 0.20413 | 0.23383 | 0.39
2024-07-03 14:24 | N2 |  |  | 20.706 | 13030.4958 | 2.2252 | 103275.3025 | 20.3695 |  | 1.1858 | 2.2344 | 20.7942 | 0.41704 | 0.2 | 0.41786 | 0.0017767 | 3 | 4.3027 | 0.0044135 | 0.2 | 0.20005 | 2.2252 | 2.2252 | 2.2252 | N2 |  |  | 174.0618 | 2.1944 | 98852.1947 | 19.821 | 1.1368 | 2.196 | 0.071894 | 0.1 | 0.10875 | 0.031975 | 3 | 4.3027 | 0.079429 | 0.1 | 0.12771 | 2.1973 |  | 2.1973 |  | N2 | 12/03/2025 09:59:18 | GR1 | 28 | 334.2865 | 2.2084 | 0.20507 | 641991.7035 | 20.0538 | 7.3879 | 2.2125 | 0.20545 | 0.18597 | 0.39 | 0.18597 | 0.005911 | 3 | 4.3027 | 0.014684 | 0.4 | 0.40027 | 0.20465 |  | 2.2085 | 10.4414 | 24.9878 | 6.6678 | 61.3157 | 1.1608 | 6.2416 | 2.2103 | 0.19742 | 92.5451 | 0.10395 | 0.22044 | 0.17092 | 1.2897 | -0.088672 | 0.074186 | 1.1953 | -0.011445 | 0.38654 | 0.029608
2024-07-03 14:25 | N2 |  |  | 20.705 | 13030.4958 | 2.2252 | 103273.1062 | 20.3695 |  | 1.1857 | 2.2345 | 20.8262 | 0.41664 | 0.2 | [8 empty] | 2.2252 |  | N2 |  |  | 166.722 | 2.1982 | 98962.9219 | 19.9165 | 1.1377 | 2.201 | 0.1252 | 0.1 | [11 empty] | N2 | 12/03/2025 10:06:44 | GR1 | 29 | 331.8531 | 2.2084 | 0.20358 | 641230.0938 | 20.0667 | 7.3788 | 2.2124 | 0.20394 | 0.18006 | 0.39
2024-07-03 14:26 | N2 |  |  | 20.706 | 13030.4958 | 2.2252 | 103273.1483 | 20.3707 |  | 1.1857 | 2.2346 | 20.793 | 0.4199 | 0.2 | [8 empty] | 2.2252 |  | N2 |  |  | 169.8678 | 2.1992 | 98888.9179 | 19.96 | 1.1367 | 2.202 | 0.12914 | 0.1 | [11 empty] | N2 | 12/03/2025 10:13:31 | GR1 | 30 | 334.6743 | 2.2085 | 0.20532 | 640581.1049 | 20.0867 | 7.3708 | 2.2128 | 0.20571 | 0.19188 | 0.39
2024-07-03 14:16 | N2 |  |  | 15.5715 | 13030.4958 | 2.926 | 104461.5511 | 20.3636 |  | 1.1994 | 2.9381 | 15.3906 | 0.41506 | 0.2 | 0.40995 | 0.0076254 | 3 | 4.3027 | 0.018942 | 0.2 | 0.2009 | 2.9264 | 2.926 | 2.9264 | N2 |  |  | 132.1194 | 3.0056 | 98853.974 | 19.8185 | 1.1369 | 3.01 | 0.14694 | 0.1 | 0.2226 | 0.070168 | 3 | 4.3027 | 0.17431 | 0.1 | 0.20096 | 3.0064 |  | 3.0064 |  | N2 | 12/03/2025 09:18:15 | GR1 | 23 | 248.0243 | 2.9919 | 0.20613 | 602129.5537 | 20.397 | 6.9204 | 2.9962 | 0.20642 | 0.14423 | 0.39 | 0.023069 | 0.10493 | 3 | 4.3027 | 0.26066 | 0.4 | 0.47744 | 0.20604 |  | 2.9897 | 10.1576 | 24.7777 | 5.5122 | 24.7627 | 0.1012 | 4.387 | 2.9742 | 0.29245 | 53.9275 | 0.13617 | 0.1175 | 0.1477 | 0.79554 | -0.069849 | 0.14778 | 0.47264 | -0.26938 | 0.4576 | 0.58868
2024-07-03 14:17 | N2 |  |  | 15.568 | 13030.4958 | 2.9267 | 104456.6169 | 20.3658 |  | 1.1993 | 2.9384 | 15.5884 | 0.40119 | 0.2 | [8 empty] | 2.9267 |  | N2 |  |  | 122.682 | 3.0058 | 98897.6806 | 19.958 | 1.1368 | 3.0143 | 0.28554 | 0.1 | [11 empty] | N2 | 12/03/2025 09:28:58 | GR1 | 24 | 248.2045 | 2.9892 | 0.20609 | 602129.9819 | 20.0795 | 6.928 | 2.9881 | 0.20602 | -0.037505 | 0.39
2024-07-03 14:18 | N2 |  |  | 15.5695 | 13030.4958 | 2.9265 | 104455.6788 | 20.3678 |  | 1.1993 | 2.9386 | 15.5626 | 0.41361 | 0.2 | [8 empty] | 2.9265 |  | N2 |  |  | 122.682 | 3.0079 | 98957.3561 | 19.9065 | 1.1377 | 3.0149 | 0.23532 | 0.1 | [11 empty] | N2 | 12/03/2025 09:34:17 | GR1 | 25 | 248.0477 | 2.9882 | 0.20589 | 602043.2698 | 19.9277 | 6.9306 | 2.987 | 0.20581 | -0.03752 | 0.39
[25 empty] | N2 |  |  | 66.0594 | 6.016 | 98866.5849 | 19.816 | 1.137 | 6.0201 | 0.068032 | 0.1 | 0.19765 | 0.17927 | 3 | 4.3027 | 0.44533 | 0.1 | 0.45642 | 6.0333 |  | 6.0333 |  | N2 | 07/03/2025 08:43:43 | GR3 | 10 | 180.5462 | 6.1074 | 0.3063 | 1097367.6602 | 19.7301 | 12.6542 | 6.1095 | 0.3064 | 0.034761 | 0.39 | 0.026522 | 0.024973 | 3 | 4.3027 | 0.062037 | 0.4 | 0.40478 | 0.30667 |  | 6.1096 |  |  | 0.94876 | 4.8003 | 0.16187 | 6.1032 | 6.0715 | 0.10186 | 10.9035 | 0.30284 |  |  |  | 0.095786 | 0.34148 | 0.28051 | -0.075338 | 0.26858 | 0.28051
[25 empty] | N2 |  |  | 61.8648 | 6.0385 | 98911.0832 | 19.956 | 1.137 | 6.0628 | 0.40223 | 0.1 | [11 empty] | N2 | 07/03/2025 08:49:00 | GR3 | 11 | 180.6496 | 6.1117 | 0.30669 | 1097060.7637 | 20.0664 | 12.6358 | 6.1145 | 0.30683 | 0.046335 | 0.39
[25 empty] | N2 |  |  | 63.9624 | 6.0453 | 98968.2553 | 19.897 | 1.1379 | 6.0527 | 0.12268 | 0.1 | [11 empty] | N2 | 07/03/2025 08:55:28 | GR3 | 12 | 180.8985 | 6.1098 | 0.30702 | 1095854.8156 | 19.9391 | 12.6275 | 6.1097 | 0.30701 | -0.0015295 | 0.39
[25 empty] | N2 |  |  | 67.1076 | 10.0292 | 98895.1158 | 19.8155 | 1.1374 | 10.041 | 0.11796 | 0.1 | 0.17084 | 0.047334 | 3 | 4.3027 | 0.11758 | 0.1 | 0.15436 | 10.0447 |  | 10.0447 |  | N2 | 07/03/2025 08:10:40 | GR3 | 4 | 183.031 | 10.0096 | 0.50891 | 1095337.2153 | 20.236 | 12.6084 | 9.9947 | 0.50815 | -0.14883 | 0.39 | -0.14857 | 0.0015182 | 3 | 4.3027 | 0.0037713 | 0.4 | 0.40002 | 0.50499 |  | 10.0085 |  |  | 7.1702 | 41.9712 | -0.9285 | 6.2494 | 10.0266 | 0.12944 | 48.2207 | 0.14401 |  |  |  | 0.041396 | 0.055568 | 0.74495 | -0.27801 | 0.3732 | 0.74495
[25 empty] | N2 |  |  | 61.8648 | 10.0501 | 98986.7713 | 19.8865 | 1.1381 | 10.0711 | 0.20925 | 0.1 | [11 empty] | N2 | 07/03/2025 08:14:27 | GR3 | 5 | 181.0308 | 10.0082 | 0.50327 | 1095471.8222 | 20.1201 | 12.615 | 9.9932 | 0.50252 | -0.14995 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 10.0549 | 98942.7973 | 19.955 | 1.1374 | 10.0735 | 0.1853 | 0.1 | [11 empty] | N2 | 07/03/2025 08:18:32 | GR3 | 6 | 180.866 | 10.0077 | 0.50279 | 1095579.0629 | 20.0638 | 12.6188 | 9.993 | 0.50205 | -0.14694 | 0.39
[25 empty] | N2 |  |  | 66.0594 | 14.9552 | 98946.5151 | 19.8145 | 1.1379 | 14.9799 | 0.16501 | 0.1 | 0.16518 | 0.030862 | 3 | 4.3027 | 0.076664 | 0.1 | 0.12601 | 14.975 |  | 14.975 |  | N2 | 06/03/2025 12:44:57 | GR4 | 108 | 180.3878 | 15.0278 | 0.75301 | 1101176.1675 | 19.8855 | 12.6913 | 14.973 | 0.75026 | -0.36462 | 0.39 | -0.36971 | 0.01 | 3 | 4.3027 | 0.024842 | 0.4 | 0.40077 | 0.75403 |  | 15.0334 |  |  | 10.4034 | 62.9825 | -2.3018 | 6.226 | 15.0042 | 0.11706 | 69.2085 | 0.1202 |  |  |  | 0.048118 | 0.037793 | 1.2732 | -0.48677 | 0.38232 | 1.2732
[25 empty] | N2 |  |  | 61.8648 | 14.9843 | 98994.8741 | 19.952 | 1.138 | 15.0044 | 0.1344 | 0.1 | [11 empty] | N2 | 06/03/2025 12:48:23 | GR4 | 109 | 180.9092 | 15.0338 | 0.75549 | 1100483.778 | 20.0766 | 12.6748 | 14.9792 | 0.75274 | -0.36327 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 14.9856 | 99039.5883 | 19.8785 | 1.1388 | 15.015 | 0.19612 | 0.1 | [11 empty] | N2 | 06/03/2025 12:52:24 | GR4 | 110 | 180.4028 | 15.0384 | 0.7536 | 1099474.3394 | 20.2419 | 12.6558 | 14.9811 | 0.75073 | -0.38123 | 0.39
[25 empty] | N2 |  |  | 63.9624 | 19.8909 | 99019.8068 | 19.8165 | 1.1388 | 19.9027 | 0.059206 | 0.1 | 0.11218 | 0.05065 | 3 | 4.3027 | 0.12582 | 0.1 | 0.16072 | 19.8993 |  | 19.8993 |  | N2 | 06/03/2025 13:40:10 | GR4 | 117 | 180.4082 | 20.0319 | 1.0039 | 1102077.9087 | 20.2245 | 12.6821 | 19.9268 | 0.9986 | -0.52485 | 0.39 | -0.48387 | 0.036008 | 3 | 4.3027 | 0.08945 | 0.4 | 0.40988 | 1.0089 |  | 20.0246 |  |  | 4.3426 | 38.7128 | -2.8802 | 5.9523 | 19.962 | 0.032743 | 44.6651 | 0.14963 |  |  |  | 0.079432 | 0.058672 | 1.3538 | -0.51661 | 0.38159 | 1.3538
[25 empty] | N2 |  |  | 61.8648 | 19.8975 | 99110.7604 | 19.8775 | 1.1396 | 19.9294 | 0.16013 | 0.1 | [11 empty] | N2 | 06/03/2025 13:43:39 | GR4 | 118 | 183.0319 | 20.0205 | 1.0179 | 1101400.9286 | 20.1732 | 12.6811 | 19.9266 | 1.0131 | -0.46946 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 19.9096 | 99065.7303 | 19.949 | 1.1388 | 19.9329 | 0.11719 | 0.1 | [11 empty] | N2 | 06/03/2025 13:48:09 | GR4 | 119 | 180.6742 | 20.0214 | 1.0048 | 1101202.9508 | 20.1813 | 12.6785 | 19.9298 | 1.0002 | -0.4573 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 25.0465 | 99168.6621 | 19.9755 | 1.1399 | 24.9939 | -0.20985 | 0.1 | -0.059054 | 0.14032 | 3 | 4.3027 | 0.34857 | 0.1 | 0.36263 | 25.0584 |  | 25.0584 |  | N2 | 06/03/2025 15:40:58 | GR5 | 126 | 181.4139 | 24.51 | 1.2351 | 1096966.1861 | 20.3333 | 12.6228 | 24.4063 | 1.2299 | -0.42319 | 0.39 | -0.41209 | 0.0096142 | 3 | 4.3027 | 0.023883 | 0.4 | 0.40071 | 1.2328 |  | 24.5056 |  |  | -0.44907 | 7.6044 | -2.5664 | 6.2278 | 24.782 | -0.218 | 13.8322 | 0.26888 |  |  |  | 0.15895 | 0.24333 | 0.65324 | -0.19408 | 0.29711 | 0.65324
[25 empty] | N2 |  |  | 66.0594 | 25.0572 | 99116.194 | 19.851 | 1.1398 | 25.0484 | -0.034997 | 0.1 | [11 empty] | N2 | 06/03/2025 15:45:36 | GR5 | 127 | 180.6382 | 24.5033 | 1.2295 | 1095074.0799 | 20.2887 | 12.603 | 24.4037 | 1.2245 | -0.40653 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 25.0716 | 99200.2052 | 19.9 | 1.1405 | 25.0885 | 0.067681 | 0.1 | [11 empty] | N2 | 06/03/2025 15:49:03 | GR5 | 128 | 181.2614 | 24.5034 | 1.2338 | 1096431.6781 | 20.2865 | 12.6188 | 24.4038 | 1.2287 | -0.40654 | 0.39
[25 empty] | N2 |  |  | 61.8648 | 30.016 | 99297.6488 | 19.9455 | 1.1415 | 30.0123 | -0.012217 | 0.1 | -0.09405 | 0.13068 | 3 | 4.3027 | 0.32464 | 0.1 | 0.33969 | 30.0176 |  | 30.0176 |  | N2 | 06/03/2025 15:07:15 | GR5 | 120 | 180.4787 | 29.0291 | 1.4553 | 1119854.504 | 20.3848 | 12.8845 | 28.9097 | 1.4493 | -0.41138 | 0.39 | -0.41568 | 0.0078815 | 3 | 4.3027 | 0.019579 | 0.4 | 0.40048 | 1.4562 |  | 29.0129 |  |  | -0.81507 | 8.6663 | -2.5918 | 6.2351 | 29.5153 | -0.22863 | 14.9014 | 0.25905 |  |  |  | 0.13458 | 0.21973 | 0.61246 | -0.18705 | 0.30541 | 0.61246
[25 empty] | N2 |  |  | 61.8648 | 30.0175 | 99301.3204 | 19.9395 | 1.1415 | 30.01 | -0.025167 | 0.1 | [11 empty] | N2 | 06/03/2025 15:10:44 | GR5 | 121 | 180.9908 | 29.0167 | 1.4588 | 1119931.0096 | 20.264 | 12.8908 | 28.8974 | 1.4528 | -0.41088 | 0.39
[25 empty] | N2 |  |  | 62.9136 | 30.0193 | 99225.0975 | 19.9305 | 1.1407 | 29.9458 | -0.24477 | 0.1 | [11 empty] | N2 | 06/03/2025 15:14:41 | GR5 | 122 | 180.5909 | 28.993 | 1.4544 | 1121857.5152 | 19.8828 | 12.9303 | 28.8699 | 1.4482 | -0.42478 | 0.39
Notes: E% in this case is based on Totalized Volume; ambient temperature ranged from 292.6 K to 292.88 K; Ambient pressure ranged from 99190 Pa to 100600 Pa (measurements taken in different days)
| PGM meter | [23 empty] | PGM meter | [21 empty] | PGM meter | [23 empty] | VSL |  | INRIM |  | NEL | [5 empty] | VSL |  |  | INRIM |  |  | NEL
Date and time | Fluid | TestPoint Timestamp | Test Point | Reference time | Reference volume | Q_Reference (corrected to MUT conditions) | p_abs_MUT | T_MUT | rho_MUT | rho_MUT | Q_MUT | MUT time | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average | Fluid | TestPoint Timestamp | Test Point | Test time | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | Q_MUT | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average | En (VSL-INRIM) | Fluid | TestPoint Timestamp | Nozzle | Test Point | Test time | Q_Reference | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | Q_MUT | Q_MUT | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(Vol) average | Total uncertainty | Q Ref average | Q Ref | Q Ref average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | Q ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
| - | - | - | s | ml | m3/h | Pa | °C | kg/m³ | kg/m3 | m3 | s | % | % | % |  |  |  | % | % | % | m3/h | m3/h | m3/h | - | - | - | s | m3/h | Pa | °C | kg/m³ | m3/h | % | % | % |  |  |  | % | % | % | m3 | m3/h | m3/h |  | - | - | - | - | s | m3/h | m3 | Pa | °C | kg/m³ | m3/h | m3 | % | % | % |  |  |  | % | % | % | m3 | m3/h | m3/h | [6 empty] | m3/h | % |  |  |  |  | % |  |  | % |  |  | %
2024-07-04 10:41 | CO2 |  |  | 73.596 | 2064.1923 | 0.10076 | 101284.7624 | 20.2596 |  | 1.1633 | 0.10004 | 73.8722 | -0.71077 | 0.2 | -0.77333 | 0.10021 | 3 | 4.3027 | 0.24893 | 0.2 | 0.31932 | 0.10077 | 0.10076 | 0.10077 | CO2 |  |  | 1901.088 | 0.096423 | 98760.641 | 19.7905 | 1.7845 | 0.094182 | -2.3239 | 0.1 | -0.59352 | 2.3861 | 3 | 4.3027 | 5.9273 | 0.1 | 5.9282 | 0.097933 |  | 0.097933 | 0.030288 | CO2 | [23 empty] | -7.5842 | 9.8072 | -0.016888 | 0.028455 |  |  | 0.09935 | -0.77281 | 9.8356 | 0.31886 | -0.0005202 | 0.017175 | 0.030288 | 0.17929 | 5.9196 | 0.030288
2024-07-04 10:42 | CO2 |  |  | 73.588 | 2064.1923 | 0.10076 | 101275.5501 | 20.265 |  | 1.1632 | 0.10004 | 73.8748 | -0.72031 | 0.2 | [8 empty] | 0.10076 |  | CO2 |  |  | 1851.8022 | 0.097885 | 98778.6566 | 19.9055 | 1.7841 | 0.096333 | -1.5852 | 0.1 | [11 empty] | CO2
2024-07-04 10:44 | CO2 |  |  | 73.5745 | 2064.1923 | 0.10078 | 101280.4642 | 20.269 |  | 1.1632 | 0.099884 | 73.2558 | -0.88891 | 0.2 | [8 empty] | 0.10078 |  | CO2 |  |  | 1898.991 | 0.09949 | 98753.3698 | 19.9255 | 1.7835 | 0.10161 | 2.1285 | 0.1 | [11 empty] | CO2
2024-07-04 10:24 | CO2 |  |  | 41.8445 | 3474.9427 | 0.29818 | 101325.8403 | 20.2195 |  | 1.164 | 0.29857 | 41.9072 | 0.12988 | 0.2 | 0.12326 | 0.0057305 | 3 | 4.3027 | 0.014235 | 0.2 | 0.20051 | 0.29824 | 0.29818 | 0.29824 | CO2 |  |  | 600.8238 | 0.30535 | 98748.0952 | 19.812 | 1.7841 | 0.30434 | -0.33268 | 0.1 | -0.25576 | 0.067007 | 3 | 4.3027 | 0.16645 | 0.1 | 0.19418 | 0.30663 |  | 0.30663 | 1.3579 | CO2 | [23 empty] | 3.066 | 24.874 | -6.7829 | 26.5202 |  |  | 0.30244 | -0.072322 | 51.3942 | 0.13949 | 0.19558 | 0.14403 | 1.3579 | -0.18344 | 0.13509 | 1.3579
2024-07-04 10:25 | CO2 |  |  | 41.831 | 3474.9427 | 0.29828 | 101325.6179 | 20.2207 |  | 1.164 | 0.29864 | 41.8976 | 0.11989 | 0.2 | [8 empty] | 0.29828 |  | CO2 |  |  | 591.3864 | 0.30724 | 98798.1612 | 19.8685 | 1.7847 | 0.3066 | -0.21003 | 0.1 | [11 empty] | CO2
2024-07-04 10:26 | CO2 |  |  | 41.8315 | 3474.9427 | 0.29827 | 101325.6937 | 20.2229 |  | 1.164 | 0.29863 | 41.8992 | 0.12001 | 0.2 | [8 empty] | 0.29827 |  | CO2 |  |  | 591.3864 | 0.30729 | 98847.3028 | 19.9235 | 1.7853 | 0.3066 | -0.22458 | 0.1 | [11 empty] | CO2
2024-07-04 09:31 | CO2 |  |  | 63.0905 | 10615.1509 | 0.60312 | 101490.6557 | 20.2051 |  | 1.1659 | 0.60517 | 63.1148 | 0.33988 | 0.2 | 0.34181 | 0.0042794 | 3 | 4.3027 | 0.010631 | 0.2 | 0.20028 | 0.6032 | 0.60312 | 0.6032 | CO2 |  |  | 305.1324 | 0.59541 | 98743.1877 | 19.815 | 1.784 | 0.59758 | 0.3654 | 0.1 | -0.034266 | 0.4126 | 3 | 4.3027 | 1.025 | 0.1 | 1.0298 | 0.59946 |  | 0.59946 | 0.35846 | CO2 | 14/03/2025 11:23:16 | RS6 | 35 | 1180.0644 | 0.61525 | 0.20168 | 626141.7572 | 19.5466 | 11.7197 | 0.62279 | 0.20415 | 1.2246 | 0.39 | 1.2733 | 0.047603 | 3 | 4.3027 | 0.11825 | 0.4 | 0.41711 | 0.20139 |  | 0.61449 | 8.5211 | 24.9296 | -0.03231 | 0.94291 | 7.3184 | 5.7477 | 0.60572 | 0.49991 | 31.6202 | 0.17784 | -0.1581 | 0.092128 | 1.7161 | -0.53417 | 1.0144 | 0.52661 | 0.77337 | 0.3773 | 2.0497
2024-07-04 09:32 | CO2 |  |  | 63.08 | 10615.1509 | 0.60322 | 101492.7559 | 20.2075 |  | 1.1659 | 0.60527 | 62.9836 | 0.33883 | 0.2 | [8 empty] | 0.60322 |  | CO2 |  |  | 307.2294 | 0.59999 | 98842.3061 | 19.92 | 1.7852 | 0.59993 | -0.0095179 | 0.1 | [11 empty] | CO2 | 14/03/2025 11:43:40 | RS6 | 36 | 1180.7653 | 0.61427 | 0.20148 | 625094.2161 | 19.3371 | 11.7087 | 0.62211 | 0.20404 | 1.2754 | 0.39
2024-07-04 09:34 | CO2 |  |  | 63.0765 | 10615.1509 | 0.60326 | 101499.7096 | 20.2083 |  | 1.166 | 0.60535 | 63.0958 | 0.34671 | 0.2 | [8 empty] | 0.60326 |  | CO2 |  |  | 305.1324 | 0.60299 | 98809.1424 | 19.835 | 1.7851 | 0.60022 | -0.45869 | 0.1 | [11 empty] | CO2 | 14/03/2025 12:04:43 | RS6 | 37 | 1178.7492 | 0.61395 | 0.20103 | 622338.6154 | 19.2671 | 11.6584 | 0.62205 | 0.20368 | 1.3198 | 0.39
2024-07-04 09:19 | CO2 |  |  | 46.723 | 13030.4958 | 0.99643 | 101808.0827 | 20.1904 |  | 1.1696 | 0.99997 | 46.7574 | 0.35512 | 0.2 | 0.35557 | 0.0040142 | 3 | 4.3027 | 0.0099719 | 0.2 | 0.20025 | 0.99634 | 0.99643 | 0.99634 | CO2 |  |  | 204.4704 | 1.0041 | 98838.9732 | 19.917 | 1.7851 | 1.0052 | 0.11232 | 0.1 | 0.032952 | 0.09432 | 3 | 4.3027 | 0.2343 | 0.1 | 0.25475 | 1.0074 |  | 1.0074 | 0.99564 | CO2 | 14/03/2025 10:45:48 | RS6 | 32 | 688.2906 | 1.0557 | 0.20183 | 595309.3433 | 20.3448 | 11.0893 | 1.0683 | 0.20425 | 1.1965 | 0.39 | 1.1001 | 0.085091 | 3 | 4.3027 | 0.21138 | 0.4 | 0.45242 | 0.20201 |  | 1.0554 | 8.8673 | 24.938 | 0.50775 | 15.4088 | 5.3748 | 4.8857 | 1.0197 | 0.32609 | 45.2325 | 0.14869 | 0.029483 | 0.13413 | 0.2198 | -0.29314 | 0.20686 | 1.4171 | 0.77403 | 0.42728 | 1.8115
2024-07-04 09:20 | CO2 |  |  | 46.729 | 13030.4958 | 0.99632 | 101807.7225 | 20.1929 |  | 1.1696 | 0.99982 | 46.6912 | 0.35181 | 0.2 | [8 empty] | 0.99632 |  | CO2 |  |  | 183.4992 | 1.0059 | 98813.0871 | 19.807 | 1.7853 | 1.0065 | 0.057851 | 0.1 | [11 empty] | CO2 | 14/03/2025 10:58:01 | RS6 | 33 | 689.2039 | 1.0555 | 0.20207 | 596772.6728 | 20.3292 | 11.1181 | 1.0668 | 0.20423 | 1.0687 | 0.39
2024-07-04 09:22 | CO2 |  |  | 46.732 | 13030.4958 | 0.99626 | 101804.9689 | 20.1953 |  | 1.1696 | 0.99984 | 46.7634 | 0.3598 | 0.2 | [8 empty] | 0.99626 |  | CO2 |  |  | 191.8878 | 1.0124 | 98729.8096 | 20.303 | 1.7808 | 1.0116 | -0.071319 | 0.1 | [11 empty] | CO2 | 14/03/2025 11:10:14 | RS6 | 34 | 689.5946 | 1.0551 | 0.20211 | 597846.2134 | 20.2661 | 11.1414 | 1.0661 | 0.20421 | 1.0352 | 0.39
2024-07-04 09:07 | CO2 |  |  | 20.471 | 13030.4958 | 2.2316 | 103756.9455 | 20.1832 |  | 1.192 | 2.2401 | 20.2848 | 0.37867 | 0.2 | 0.36426 | 0.014959 | 3 | 4.3027 | 0.03716 | 0.2 | 0.20342 | 2.2319 | 2.2316 | 2.2319 | CO2 |  |  | 174.0618 | 2.1919 | 98695.9708 | 20.266 | 1.7804 | 2.1948 | 0.13275 | 0.1 | 0.082367 | 0.06633 | 3 | 4.3027 | 0.16477 | 0.1 | 0.19274 | 2.1943 |  | 2.1943 | 1.0059 | CO2 | 13/03/2025 08:10:39 | GR1 | 1 | 335.0824 | 2.1886 | 0.20371 | 585941.4158 | 19.838 | 10.9299 | 2.2094 | 0.20565 | 0.95201 | 0.39 | 0.99014 | 0.07697 | 3 | 4.3027 | 0.1912 | 0.4 | 0.44335 | 0.2033 |  | 2.1886 | 8.8026 | 24.1658 | 2.2171 | 26.9178 | 5.0374 | 5.0875 | 2.2049 | 0.28586 | 56.1711 | 0.13343 | 0.078398 | 0.15355 | 0.51056 | -0.20349 | 0.13909 | 1.463 | 0.70428 | 0.4228 | 1.6658
2024-07-04 09:09 | CO2 |  |  | 20.47 | 13030.4958 | 2.2318 | 103761.3142 | 20.1832 |  | 1.1921 | 2.2399 | 20.4492 | 0.3653 | 0.2 | [8 empty] | 2.2318 |  | CO2 |  |  | 170.916 | 2.1934 | 98839.3446 | 19.916 | 1.7852 | 2.1958 | 0.10713 | 0.1 | [11 empty] | CO2 | 13/03/2025 08:16:49 | GR1 | 2 | 333.3722 | 2.1892 | 0.20273 | 589367.3026 | 19.907 | 10.9931 | 2.2098 | 0.20463 | 0.93969 | 0.39
2024-07-04 09:10 | CO2 |  |  | 20.4645 | 13030.4958 | 2.2324 | 103757.2102 | 20.1829 |  | 1.1921 | 2.2402 | 20.6104 | 0.34881 | 0.2 | [8 empty] | 2.2324 |  | CO2 |  |  | 176.1594 | 2.1975 | 98818.3527 | 19.7715 | 1.7857 | 2.1977 | 0.007218 | 0.1 | [11 empty] | CO2 | 13/03/2025 08:22:45 | GR1 | 3 | 334.7473 | 2.188 | 0.20346 | 588524.0422 | 20.0091 | 10.9726 | 2.2116 | 0.20565 | 1.0787 | 0.39
[25 empty] | CO2 |  |  | 123.7308 | 3.0086 | 98840.6434 | 19.9135 | 1.7852 | 3.0101 | 0.049273 | 0.1 | 0.095692 | 0.052892 | 3 | 4.3027 | 0.13139 | 0.1 | 0.16512 | 3.012 |  | 3.012 |  | CO2 | 13/03/2025 11:53:29 | GR2 | 14 | 243.236 | 3.0397 | 0.20538 | 625527.4579 | 20.1102 | 11.6826 | 3.0582 | 0.20663 | 0.60741 | 0.39 | 0.591 | 0.027078 | 3 | 4.3027 | 0.067267 | 0.4 | 0.40562 | 0.20504 |  | 3.0395 |  |  | 3.5099 | 36.6792 | 3.5921 | 6.0781 | 3.0258 | 0.1661 | 42.7573 | 0.15293 |  |  |  | -0.07041 | 0.062254 | 1.131 | 0.4249 | 0.37568 | 1.131
[25 empty] | CO2 |  |  | 125.8278 | 3.0105 | 98820.4112 | 19.7405 | 1.7859 | 3.0152 | 0.15327 | 0.1 | [11 empty] | CO2 | 13/03/2025 11:57:56 | GR2 | 15 | 242.447 | 3.0388 | 0.20465 | 639545.4986 | 19.9395 | 11.9617 | 3.0572 | 0.20589 | 0.60584 | 0.39
[25 empty] | CO2 |  |  | 127.9248 | 3.0169 | 98669.8358 | 20.2515 | 1.7801 | 3.0195 | 0.084528 | 0.1 | [11 empty] | CO2 | 13/03/2025 12:02:26 | GR2 | 16 | 242.8655 | 3.0401 | 0.20509 | 646783.5886 | 19.9623 | 12.1011 | 3.0571 | 0.20624 | 0.55974 | 0.39
[25 empty] | CO2 |  |  | 66.0594 | 6.0319 | 98853.0722 | 19.9105 | 1.7854 | 6.0334 | 0.023604 | 0.1 | 0.010626 | 0.041625 | 3 | 4.3027 | 0.1034 | 0.1 | 0.14385 | 6.0402 |  | 6.0402 |  | CO2 | 13/03/2025 14:11:05 | GR3 | 23 | 180.625 | 5.9955 | 0.30082 | 650234.721 | 20.1892 | 12.1575 | 6.0355 | 0.30283 | 0.66824 | 0.39 | 0.54181 | 0.1109 | 3 | 4.3027 | 0.27548 | 0.4 | 0.48568 | 0.30463 |  | 5.9965 |  |  | 0.51355 | 48.3278 | 2.2969 | 4.2393 | 6.0184 | 0.053464 | 52.5671 | 0.13793 |  |  |  | -0.042838 | 0.04085 | 1.0487 | 0.48835 | 0.46569 | 1.0487
[25 empty] | CO2 |  |  | 61.8648 | 6.0425 | 98659.6588 | 20.2375 | 1.78 | 6.0403 | -0.035941 | 0.1 | [11 empty] | CO2 | 13/03/2025 14:17:24 | GR3 | 24 | 187.0399 | 5.9977 | 0.31161 | 661467.8154 | 20.0828 | 12.3805 | 6.0253 | 0.31305 | 0.461 | 0.39
[25 empty] | CO2 |  |  | 62.9136 | 6.0463 | 98840.2804 | 19.6695 | 1.7867 | 6.0489 | 0.044217 | 0.1 | [11 empty] | CO2 | 13/03/2025 14:21:01 | GR3 | 25 | 180.99 | 5.9964 | 0.30147 | 664702.288 | 20.1271 | 12.4412 | 6.0262 | 0.30297 | 0.4962 | 0.39
[25 empty] | CO2 |  |  | 61.8648 | 10.025 | 98726.6452 | 20.0465 | 1.7823 | 10.0581 | 0.33082 | 0.1 | 0.28523 | 0.18358 | 3 | 4.3027 | 0.45604 | 0.1 | 0.46687 | 10.0319 |  | 10.0319 |  | CO2 | 13/03/2025 13:58:30 | GR3 | 20 | 181.1087 | 10.1865 | 0.51246 | 504798.817 | 18.7506 | 9.4115 | 10.2123 | 0.51376 | 0.2536 | 0.39 | 0.22946 | 0.026571 | 3 | 4.3027 | 0.066007 | 0.4 | 0.40541 | 0.52563 |  | 10.1977 |  |  | 1.3086 | 4.5878 | 1.3961 | 6.0843 | 10.1148 | 0.25343 | 10.6721 | 0.30611 |  |  |  | 0.031792 | 0.35252 | 0.090186 | -0.023972 | 0.26581 | 0.090186
[25 empty] | CO2 |  |  | 61.8648 | 10.0339 | 98755.485 | 20.112 | 1.7824 | 10.0782 | 0.44172 | 0.1 | [11 empty] | CO2 | 13/03/2025 14:01:57 | GR3 | 21 | 188.196 | 10.1964 | 0.53303 | 507448.2245 | 19.1106 | 9.4496 | 10.2169 | 0.53411 | 0.20099 | 0.39
[25 empty] | CO2 |  |  | 62.9136 | 10.0368 | 98885.6227 | 19.9095 | 1.786 | 10.0451 | 0.083147 | 0.1 | [11 empty] | CO2 | 13/03/2025 14:05:50 | GR3 | 22 | 187.3654 | 10.2102 | 0.5314 | 485549.8997 | 19.5995 | 9.0141 | 10.2341 | 0.53264 | 0.23379 | 0.39
[25 empty] | CO2 |  |  | 61.8648 | 14.964 | 98835.3501 | 20.0975 | 1.784 | 14.9784 | 0.096217 | 0.1 | 0.068932 | 0.080443 | 3 | 4.3027 | 0.19983 | 0.1 | 0.22346 | 14.9741 |  | 14.9741 |  | CO2 | 14/03/2025 08:16:58 | GR4 | 26 | 187.1306 | 15.3612 | 0.79849 | 331630.735 | 20.9462 | 6.0746 | 15.4521 | 0.80321 | 0.59161 | 0.39 | 0.083963 | 0.45952 | 3 | 4.3027 | 1.1415 | 0.4 | 1.2096 | 0.77962 |  | 15.3369 |  |  | 1.3805 | 20.0271 | 0.057389 | 0.6835 | 15.1555 | 0.069428 | 20.7106 | 0.21974 |  |  |  | -0.00049607 | 0.040594 | 0.01222 | 0.014535 | 1.1894 | 0.01222
[25 empty] | CO2 |  |  | 63.9624 | 14.9719 | 98983.181 | 19.818 | 1.7883 | 14.9917 | 0.13218 | 0.1 | [11 empty] | CO2 | 14/03/2025 08:23:46 | GR4 | 27 | 180.8065 | 15.336 | 0.77024 | 399876.0479 | 20.0027 | 7.3773 | 15.3305 | 0.76996 | -0.036143 | 0.39
[25 empty] | CO2 |  |  | 61.8648 | 14.9863 | 98964.2356 | 19.689 | 1.7888 | 14.9831 | -0.021604 | 0.1 | [11 empty] | CO2 | 14/03/2025 08:33:28 | Gr4 | 28 | 181.0508 | 15.3135 | 0.77014 | 431114.5326 | 19.3972 | 7.9852 | 15.267 | 0.7678 | -0.30358 | 0.39
[25 empty] | CO2 |  |  | 52.428 | 19.8978 | 99086.8904 | 19.8375 | 1.7901 | 19.8935 | -0.021563 | 0.1 | -0.036331 | 0.11306 | 3 | 4.3027 | 0.28086 | 0.1 | 0.29813 | 19.9053 |  | 19.9053 |  | CO2
[25 empty] | CO2 |  |  | 61.8648 | 19.906 | 98932.9303 | 20.0975 | 1.7857 | 19.875 | -0.15605 | 0.1 | [11 empty] | CO2
[25 empty] | CO2 |  |  | 61.8648 | 19.9122 | 99075.1048 | 19.6445 | 1.7911 | 19.9258 | 0.068622 | 0.1 | [11 empty] | CO2
[25 empty] | CO2 |  |  | 61.8648 | 25.0504 | 99224.234 | 19.746 | 1.7931 | 25.0673 | 0.067403 | 0.1 | 0.027324 | 0.090637 | 3 | 4.3027 | 0.22515 | 0.1 | 0.24636 | 25.0592 |  | 25.0592 |  | CO2
[25 empty] | CO2 |  |  | 61.8648 | 25.0598 | 99065.8667 | 19.97 | 1.7889 | 25.0826 | 0.091011 | 0.1 | [11 empty] | CO2
[25 empty] | CO2 |  |  | 61.8648 | 25.0675 | 99212.7366 | 19.736 | 1.793 | 25.0483 | -0.076444 | 0.1 | [11 empty] | CO2
[25 empty] | CO2 |  |  | 61.8648 | 30.0139 | 99222.7603 | 19.934 | 1.792 | 30.0052 | -0.028978 | 0.1 | -0.053741 | 0.024355 | 3 | 4.3027 | 0.060502 | 0.1 | 0.11688 | 30.019 |  | 30.019 |  | CO2
[25 empty] | CO2 |  |  | 61.8648 | 30.0204 | 99382.5654 | 19.6675 | 1.7965 | 30.0041 | -0.05458 | 0.1 | [11 empty] | CO2
[25 empty] | CO2 |  |  | 61.8648 | 30.0226 | 99383.2183 | 19.709 | 1.7962 | 29.9993 | -0.077666 | 0.1 | [11 empty] | CO2
Note on G11: Ara Abdulrahman: Not corrected to MUT conditions
Note on G59: Ara Abdulrahman: Not corrected to MUT conditions

## VSL Densities
| Pgauge (Pa) | barometric pressure (Pa) | Absolute pressure (Pa) | T (°C) | density (kg/m3)
| 512.4003 | 100912 | 101424.4003 | 20.3378 | 1.1646
| 512.5354 | 100907 | 101419.5354 | 20.3361 | 1.1646
| 512.3448 | 100907 | 101419.3448 | 20.3355 | 1.1646
| 568.52 | 100935 | 101503.52 | 20.3429 | 1.1655
| 568.4343 | 100930 | 101498.4343 | 20.3421 | 1.1655
| 568.4735 | 100933 | 101501.4735 | 20.3422 | 1.1655
| 775.9324 | 100944 | 101719.9324 | 20.3485 | 1.168
| 775.6277 | 100937 | 101712.6277 | 20.3469 | 1.1679
| 775.4998 | 100947 | 101722.4998 | 20.3467 | 1.168
| 1009.0581 | 100958 | 101967.0581 | 20.3589 | 1.1708
| 1008.9877 | 100956 | 101964.9877 | 20.3571 | 1.1708
N2 | 1009.0071 | 100951 | 101960.0071 | 20.3566 | 1.1707
| 2307.3025 | 100968 | 103275.3025 | 20.3695 | 1.1858
| 2308.1062 | 100965 | 103273.1062 | 20.3695 | 1.1857
| 2308.1483 | 100965 | 103273.1483 | 20.3707 | 1.1857
| 3477.5511 | 100984 | 104461.5511 | 20.3636 | 1.1994
| 3477.6169 | 100979 | 104456.6169 | 20.3658 | 1.1993
| 3478.6788 | 100977 | 104455.6788 | 20.3678 | 1.1993
| 514.7624 | 100770 | 101284.7624 | 20.2596 | 1.1633
| 514.5501 | 100761 | 101275.5501 | 20.265 | 1.1632
| 514.4642 | 100766 | 101280.4642 | 20.269 | 1.1632
| 583.8403 | 100742 | 101325.8403 | 20.2195 | 1.164
| 583.6179 | 100742 | 101325.6179 | 20.2207 | 1.164
| 583.6937 | 100742 | 101325.6937 | 20.2229 | 1.164
CO2 | 827.6557 | 100663 | 101490.6557 | 20.2051 | 1.1659
| 827.7559 | 100665 | 101492.7559 | 20.2075 | 1.1659
| 827.7096 | 100672 | 101499.7096 | 20.2083 | 1.166
| 1152.0827 | 100656 | 101808.0827 | 20.1904 | 1.1696
| 1151.7225 | 100656 | 101807.7225 | 20.1929 | 1.1696
| 1151.9689 | 100653 | 101804.9689 | 20.1953 | 1.1696
| 3112.9455 | 100644 | 103756.9455 | 20.1832 | 1.192
| 3113.3142 | 100648 | 103761.3142 | 20.1832 | 1.1921
| 3113.2102 | 100644 | 103757.2102 | 20.1829 | 1.1921

## Sheet2
0 | 1 | 1.2
35 | 1 | 1.2
Graph
3.5 | 0 | 0.21467 | VSL
3.7 | 0 | 0.60016 | INRIM
3.95 | 0 | 0.49178 | NEL
Graph
32 | 0 | 0.21467 | VSL
33 | 0 | 0.60016 | INRIM
34 | 0 | 0.49178 | NEL

## Factory Cal Cert
Index | Flow Rate [m3/h] | Deviation [%]
1 | 30 | 0.18
2 | 20 | 0.21
3 | 14.44 | 0.21
4 | 9.98 | 0.24
5 | 7 | 0.22
6 | 4.32 | 0.22
7 | 2 | 0.19
8 | 0.99 | 0.17
9 | 0.5 | 0.1
10 | 0.24 | -0.23