## Data
[110 empty] | (-) | (%)
NEL | [29 empty] | FORCE | [29 empty] | DNV | [48 empty] | Total tes points | 35
[109 empty] | Above 1 | 3 | 8.5714
[109 empty] | Above 1.2 | 1 | 2.8571
Cp/Cv |  | 1.63 | [-] | [26 empty] | Cp/Cv |  | 1.61 | [-] | [26 empty] | Cp/Cv |  | 1.59 | [-] | [45 empty] | Between 1 and 1.2 | 2 | 5.7143
molecular mass |  | 44.01 | g/mol | [26 empty] | molecular mass |  | 44.01 | g/mol | [26 empty] | molecular mass |  | 44.01 | g/mol | [45 empty] | Below equal 1 | 32 | 91.4286
|  | CO2 | [29 empty] | CO2 | [29 empty] | CO2
Test Number | T03973 | [24 empty] | 0.35968 | [29 empty] | 0.2878 | [29 empty] | 0.27573
Coriolis Meter | [29 empty] | Coriolis Meter | [29 empty] | Coriolis Meter | [29 empty] | NEL |  | FORCE |  | DNV | [5 empty] | NEL |  |  | FORCE |  |  | DNV
Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(Mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | M ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
- | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | - | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | - | - | - | kg/s | bar | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | [6 empty] | kg/s | % |  |  |  |  | % |  |  | % |  |  | %
CO2 | 10/07/2024 10:02:09 | 4 | 7.7733 | 31.1423 | 20.0381 | 69.3521 | 7.782 | 238.7202 | 300.4667 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0022 | 7.7877 | 0.18569 | 0.35 | 0.13677 | 0.047147 | 3 | 4.3027 | 0.11712 | 0.35 | 0.36908 | 7.7723 | 403.9578 | 403.6669 | CO2 | [29 empty] | CO2 | 2025-01-13 11:51 | 1.1 | 6.7156 | 29.0817 | 24.1576 | 61.9304 | 6.7293 | 244.0429 | 298.8346 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0021 | 6.7336 | 0.26796 | 0.21831 | 0.21226 | 0.050475 | 3 | 4.3027 | 0.12539 | 0.21839 | 0.25183 | 6.7173 | 391.1718 | 390.8515 | 1.0041 | 7.3412 |  |  | 3.3471 | 15.7687 | 7.2448 | 0.18828 | 23.1099 | 0.20802 | -0.05151 | 0.30487 | 0.16896 |  |  |  | 0.023981 | 0.14193 | 0.16896
CO2 | 10/07/2024 10:11:43 | 5 | 7.7709 | 31.1538 | 20.0916 | 69.3599 | 7.7756 | 238.75 | 300.4941 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0022 | 7.7813 | 0.133 | 0.35 | [8 empty] | 403.5782 |  | CO2 | [29 empty] | CO2 | 2025-01-13 11:53 | 1.2 | 6.7177 | 29.0822 | 24.0675 | 61.9606 | 6.7268 | 243.9766 | 298.7893 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0021 | 6.7311 | 0.19928 | 0.21814 | [8 empty] | 390.8367
CO2 | 10/07/2024 10:21:47 | 6 | 7.7726 | 31.1577 | 20.103 | 69.3657 | 7.7741 | 238.7472 | 300.4999 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0022 | 7.7797 | 0.091624 | 0.35 | [8 empty] | 403.4648 |  | CO2 | [29 empty] | CO2 | 2025-01-13 11:59 | 1.3 | 6.7186 | 29.0789 | 23.9623 | 61.9965 | 6.7257 | 243.8954 | 298.7364 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0021 | 6.7299 | 0.16955 | 0.21873 | [8 empty] | 390.5461
CO2 | 10/07/2024 10:47:45 | 7 | 5.4389 | 31.3239 | 20.2379 | 69.7758 | 5.4424 | 238.6863 | 300.569 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 5.4464 | 0.13838 | 0.35 | 0.14193 | 0.038758 | 3 | 4.3027 | 0.096279 | 0.35 | 0.363 | 5.4358 | 280.7923 | 280.7249 | CO2 | 2025-05-28 |  | 5.3234 | 31.1587 | 22.1335 | 68.4426 | 5.3246 | 240.3776 | 299.6828 | 93.52 | 44.7 | 0.02235 | 1 | 0.99851 | -0.0077 | 1.0022 | 5.3286 | 0.097914 | 0.27 | 0.10805 | 0.014553 | 3 | 4.3027 | 0.036152 | 0.27 | 0.27241 | 5.3243 | 280.0678 | 280.1638 | CO2 | 2025-01-13 12:08 | 2.1 | 5.4119 | 30.3716 | 25.4197 | 64.8277 | 5.4052 | 243.7134 | 299.4682 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0022 | 5.4092 | -0.051092 | 0.2162 | 0.012679 | 0.055314 | 3 | 4.3027 | 0.13741 | 0.21612 | 0.2561 | 5.4109 | 300.1617 | 300.3136 | 1.0771 | 7.589 | 1.456 | 13.4758 | 0.19332 | 15.2469 | 5.3903 | 0.075085 | 36.3118 | 0.16595 | 0.066846 | 0.32285 | 0.20705 | 0.032963 | 0.21603 | 0.15259 | -0.062406 | 0.19506 | 0.31993
CO2 | 10/07/2024 10:58:47 | 8 | 5.4367 | 31.3038 | 20.171 | 69.7501 | 5.4384 | 238.6511 | 300.5347 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 5.4424 | 0.10507 | 0.35 | [8 empty] | 280.6894 |  | CO2 | 2025-05-28 |  | 5.3253 | 31.145 | 22.0064 | 68.441 | 5.3267 | 240.3034 | 299.6183 | 93.52 | 44.7 | 0.02235 | 1 | 0.99851 | -0.0077 | 1.0022 | 5.3307 | 0.10151 | 0.27 | [8 empty] | 280.186 |  | CO2 | 2025-01-13 12:10 | 2.2 | 5.4102 | 30.3668 | 25.396 | 64.8221 | 5.4089 | 243.6984 | 299.4563 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0022 | 5.4128 | 0.047657 | 0.2161 | [8 empty] | 300.3898
CO2 | 10/07/2024 11:08:29 | 9 | 5.4319 | 31.2662 | 19.9623 | 69.7418 | 5.4378 | 238.5125 | 300.4278 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 5.4418 | 0.18234 | 0.35 | [8 empty] | 280.6932 |  | CO2 | 2025-05-28 |  | 5.3242 | 31.1139 | 21.8455 | 68.4296 | 5.3268 | 240.208 | 299.5366 | 93.52 | 44.7 | 0.02235 | 1 | 0.99851 | -0.0077 | 1.0022 | 5.3308 | 0.12472 | 0.27 | [8 empty] | 280.2377 |  | CO2 | 2025-01-13 12:12 | 2.3 | 5.4104 | 30.3648 | 25.3844 | 64.8207 | 5.4087 | 243.6932 | 299.4505 | 93.52 | 44.7 | 0.02235 | 1 | 0.99855 | -0.0077 | 1.0022 | 5.4127 | 0.041472 | 0.21605 | [8 empty] | 300.3895
CO2 | 10/07/2024 11:20:27 | 10 | 3.1157 | 31.3674 | 20.0093 | 70.0089 | 3.1184 | 238.4435 | 300.4519 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 3.1207 | 0.15963 | 0.35 | 0.17016 | 0.01439 | 3 | 4.3027 | 0.035747 | 0.35 | 0.35182 | 3.1134 | 160.354 | 160.325 | CO2 | 2025-05-28 |  | 3.0503 | 31.0444 | 21.0395 | 68.6014 | 3.0429 | 239.6359 | 299.1272 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 3.0451 | -0.17088 | 0.27 | -0.16244 | 0.0087076 | 3 | 4.3027 | 0.021631 | 0.27 | 0.27087 | 3.0503 | 159.6799 | 159.6869 | CO2 | 2025-01-13 13:09 | 4.1 | 2.9962 | 31.6587 | 26.2066 | 67.9008 | 2.9993 | 243.0506 | 299.8626 | 93.52 | 44.7 | 0.02235 | 1 | 0.99854 | -0.0077 | 1.0023 | 3.0018 | 0.18764 | 0.22895 | 0.18183 | 0.0050437 | 3 | 4.3027 | 0.012529 | 0.22457 | 0.22492 | 2.9928 | 159.02 | 158.986 | 1.3747 | 8.079 | -2.214 | 13.6299 | 3.5943 | 19.7672 | 3.0522 | 0.066423 | 41.4762 | 0.15527 | 0.10374 | 0.3157 | 0.3286 | -0.22886 | 0.22194 | 1.0312 | 0.11541 | 0.16272 | 0.70922
CO2 | 10/07/2024 11:31:27 | 11 | 3.1124 | 31.3614 | 20.0489 | 69.9729 | 3.1159 | 238.4872 | 300.4722 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 3.1182 | 0.18656 | 0.35 | [8 empty] | 160.3094 |  | CO2 | 2025-05-28 |  | 3.0503 | 31.0301 | 20.972 | 68.6052 | 3.0431 | 239.5905 | 299.0928 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 3.0453 | -0.16296 | 0.27 | [8 empty] | 159.6845 |  | CO2 | 2025-01-13 13:11 | 4.2 | 2.992 | 31.6153 | 26.1128 | 67.8239 | 2.9949 | 243.0131 | 299.8156 | 93.52 | 44.7 | 0.02235 | 1 | 0.99854 | -0.0077 | 1.0023 | 2.9973 | 0.17854 | 0.22397 | [8 empty] | 158.9631
CO2 | 10/07/2024 11:41:21 | 12 | 3.1121 | 31.3515 | 20.039 | 69.9493 | 3.1149 | 238.49 | 300.4671 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 3.1172 | 0.16429 | 0.35 | [8 empty] | 160.3115 |  | CO2 | 2025-05-28 |  | 3.0504 | 31.016 | 20.8979 | 68.6091 | 3.0435 | 239.5451 | 299.0551 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 3.0457 | -0.15348 | 0.27 | [8 empty] | 159.6962 |  | CO2 | 2025-01-13 13:13 | 4.3 | 2.9903 | 31.5853 | 26.0211 | 67.7819 | 2.9932 | 242.9736 | 299.7697 | 93.52 | 44.7 | 0.02235 | 1 | 0.99854 | -0.0077 | 1.0023 | 2.9957 | 0.17931 | 0.22078 | [8 empty] | 158.975
CO2 | 10/07/2024 12:05:59 | 14 | 1.5654 | 31.3745 | 20.0195 | 70.0244 | 1.5663 | 238.4517 | 300.4571 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 1.5675 | 0.1313 | 0.35 | 0.12877 | 0.0064358 | 3 | 4.3027 | 0.015987 | 0.35 | 0.35036 | 1.5649 | 80.5254 | 80.5237 | CO2 | 2025-05-28 |  | 1.5251 | 31.0544 | 21.1006 | 68.6021 | 1.5213 | 239.6733 | 299.1582 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 1.5224 | -0.17992 | 0.27 | -0.18972 | 0.0094144 | 3 | 4.3027 | 0.023387 | 0.27 | 0.27101 | 1.5253 | 79.8309 | 79.8141 | CO2 | 2025-01-13 13:31 | 5.1 | 1.5692 | 31.4173 | 25.1741 | 67.7087 | 1.5701 | 242.4883 | 299.345 | 93.52 | 44.7 | 0.02235 | 1 | 0.99854 | -0.0077 | 1.0023 | 1.5714 | 0.14239 | 0.29057 | 0.13021 | 0.010823 | 3 | 4.3027 | 0.026886 | 0.26611 | 0.26747 | 1.5657 | 83.4826 | 83.4182 | 1.049 | 8.1463 | -2.5831 | 13.6153 | 1.8201 | 13.9785 | 1.552 | 0.0080023 | 35.7401 | 0.16727 | 0.12077 | 0.30786 | 0.3923 | -0.19772 | 0.21323 | 0.92728 | 0.1222 | 0.20871 | 0.58553
CO2 | 10/07/2024 12:25:14 | 15 | 1.5648 | 31.371 | 20.0488 | 70.0004 | 1.5658 | 238.4763 | 300.4721 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 1.5669 | 0.13356 | 0.35 | [8 empty] | 80.5252 |  | CO2 | 2025-05-28 |  | 1.5253 | 31.053 | 21.0736 | 68.6158 | 1.5213 | 239.6493 | 299.1445 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 1.5224 | -0.19056 | 0.27 | [8 empty] | 79.8179 |  | CO2 | 2025-01-13 13:33 | 5.2 | 1.5652 | 31.325 | 24.8899 | 67.5815 | 1.5659 | 242.3721 | 299.2024 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 1.5672 | 0.12652 | 0.2666 | [8 empty] | 83.4155
CO2 | 10/07/2024 12:45:04 | 16 | 1.5645 | 31.3675 | 20.069 | 69.9808 | 1.5652 | 238.4926 | 300.4825 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 1.5664 | 0.12146 | 0.35 | [8 empty] | 80.5205 |  | CO2 | 2025-05-28 |  | 1.5253 | 31.0466 | 21.0395 | 68.6295 | 1.5212 | 239.6252 | 299.1272 | 93.52 | 44.7 | 0.02235 | 1 | 0.9985 | -0.0077 | 1.0022 | 1.5223 | -0.19869 | 0.27 | [8 empty] | 79.7934 |  | CO2 | 2025-01-13 13:35 | 5.3 | 1.5628 | 31.2596 | 24.6099 | 67.5234 | 1.5635 | 242.216 | 299.0618 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 1.5647 | 0.12171 | 0.24116 | [8 empty] | 83.3564
CO2 | 10/07/2024 13:06:25 | 17 | 0.78427 | 31.3698 | 20.0135 | 70.0136 | 0.78452 | 238.4435 | 300.4541 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 0.78511 | 0.10623 | 0.35 | 0.1049 | 0.015786 | 3 | 4.3027 | 0.039214 | 0.35 | 0.35219 | 0.78419 | 40.3391 | 40.3388 | CO2 | [29 empty] | CO2 | 2025-01-13 14:11 | 8.1 | 0.64659 | 31.5038 | 24.3979 | 68.2593 | 0.64663 | 241.8116 | 298.9553 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.64714 | 0.085787 | 0.22625 | 0.049888 | 0.031374 | 3 | 4.3027 | 0.077938 | 0.22621 | 0.23926 | 0.64657 | 34.1032 | 34.0984 | 0.84572 | 8.0621 |  |  | 0.87147 | 17.4686 | 0.71538 | 0.06726 | 25.5306 | 0.19791 | 0.037641 | 0.29132 | 0.12921 |  |  |  | -0.017372 | 0.13445 | 0.12921
CO2 | 10/07/2024 13:25:38 | 18 | 0.7843 | 31.3583 | 19.9591 | 70.0069 | 0.78441 | 238.4135 | 300.4262 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 0.78499 | 0.088493 | 0.35 | [8 empty] | 40.3371 |  | CO2 | [29 empty] | CO2 | 2025-01-13 14:14 | 8.2 | 0.64666 | 31.5089 | 24.4598 | 68.2488 | 0.64638 | 241.8482 | 298.9864 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.6469 | 0.036157 | 0.22604 | [8 empty] | 34.0955
CO2 | 10/07/2024 13:45:59 | 19 | 0.78401 | 31.3518 | 19.94 | 69.9973 | 0.78436 | 238.4081 | 300.4164 | 93.52 | 44.7 | 0.02235 | 1 | 0.99849 | -0.0077 | 1.0023 | 0.78495 | 0.11998 | 0.35 | [8 empty] | 40.3403 |  | CO2 | [29 empty] | CO2 | 2025-01-13 14:16 | 8.3 | 0.64645 | 31.5055 | 24.5037 | 68.2185 | 0.64611 | 241.8793 | 299.0085 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.64663 | 0.02772 | 0.22634 | [8 empty] | 34.0964
CO2 | 10/07/2024 14:07:38 | 20 | 0.39745 | 31.3533 | 19.8701 | 70.0349 | 0.39816 | 238.3508 | 300.3806 | 93.52 | 44.7 | 0.02235 | 1 | 0.99848 | -0.0077 | 1.0023 | 0.39846 | 0.25372 | 0.35 | 0.31083 | 0.049576 | 3 | 4.3027 | 0.12315 | 0.35 | 0.37103 | 0.39579 | 20.4667 | 20.3858 | CO2 | [29 empty] | CO2 | 2025-01-13 14:20 | 9.1 | 0.4633 | 31.5268 | 24.5128 | 68.2669 | 0.46287 | 241.867 | 299.0131 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.46324 | -0.014107 | 0.22297 | -0.038328 | 0.020992 | 3 | 4.3027 | 0.052146 | 0.22326 | 0.22927 | 0.46332 | 24.4089 | 24.4012 | 2.2578 | 7.2639 |  |  | -0.72915 | 19.0237 | 0.42956 | 0.058152 | 26.2877 | 0.19504 | 0.25268 | 0.31564 | 0.80053 |  |  |  | -0.09648 | 0.12052 | 0.80053
CO2 | 10/07/2024 14:26:34 | 21 | 0.39542 | 31.3645 | 19.868 | 70.0684 | 0.39648 | 238.3397 | 300.3795 | 93.52 | 44.7 | 0.02235 | 1 | 0.99848 | -0.0077 | 1.0023 | 0.39677 | 0.34286 | 0.35 | [8 empty] | 20.3704 |  | CO2 | [29 empty] | CO2 | 2025-01-13 14:23 | 9.2 | 0.46338 | 31.5315 | 24.4954 | 68.2836 | 0.46277 | 241.8592 | 299.0043 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.46314 | -0.051237 | 0.2232 | [8 empty] | 24.398
CO2 | 10/07/2024 14:45:47 | 22 | 0.39452 | 31.3755 | 19.9154 | 70.0771 | 0.39555 | 238.3587 | 300.4038 | 93.52 | 44.7 | 0.02235 | 1 | 0.99848 | -0.0077 | 1.0023 | 0.39584 | 0.33589 | 0.35 | [8 empty] | 20.3202 |  | CO2 | [29 empty] | CO2 | 2025-01-13 14:25 | 9.3 | 0.46327 | 31.5254 | 24.4782 | 68.2721 | 0.46267 | 241.8436 | 298.9957 | 93.52 | 44.7 | 0.02235 | 1 | 0.99853 | -0.0077 | 1.0023 | 0.46304 | -0.049641 | 0.22362 | [8 empty] | 24.3967
Cp/Cv |  | 1.47 | [-] | [26 empty] | Cp/Cv |  | 1.46 | [-] | [26 empty] | Cp/Cv |  | 1.48 | [-]
molecular mass |  | 44.01 | g/mol | [26 empty] | molecular mass |  | 44.01 | g/mol | [26 empty] | molecular mass |  | 44.01 | g/mol
|  | CO2 | [29 empty] | CO2 | [29 empty] | CO2
Test Number | T03981
Coriolis Meter | [29 empty] | Coriolis Meter | [29 empty] | Coriolis Meter | [29 empty] | NEL |  | FORCE |  | DNV | [5 empty] | NEL |  |  | FORCE |  |  | DNV
Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(Vol) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | M ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
- | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | - | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | - | - | - | kg/s | bar | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h | [6 empty] | kg/s | % |  |  |  |  | % |  |  | % |  |  | %
CO2 | 11/07/2024 09:49:27 | 2 | 4.8673 | 21.1822 | 20.0634 | 43.5361 | 4.8776 | 248.9265 | 285.3513 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 4.878 | 0.21913 | 0.35 | 0.23753 | 0.040603 | 3 | 4.3027 | 0.10086 | 0.35 | 0.36424 | 4.8687 | 403.3284 | 403.3684 | CO2 | [29 empty] | CO2 | 2025-01-14 09:03 | 11.1 | 4.6299 | 20.9055 | 20.4316 | 42.7293 | 4.6306 | 249.419 | 286.5 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 4.6309 | 0.021913 | 0.26005 | -0.04571 | 0.231 | 3 | 4.3027 | 0.57383 | 0.23662 | 0.62071 | 4.6371 | 390.1357 | 389.8413 | 1.7903 | 7.5373 |  |  | -0.11864 | 2.5955 | 4.7529 | 0.16498 | 10.1328 | 0.31415 | 0.072552 | 0.18435 | 0.39356 |  |  |  | -0.21069 | 0.53534 | 0.39356
CO2 | 11/07/2024 09:58:57 | 3 | 4.8691 | 21.186 | 20.0074 | 43.5576 | 4.8793 | 248.8851 | 285.3241 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 4.8797 | 0.20938 | 0.35 | [8 empty] | 403.268 |  | CO2 | [29 empty] | CO2 | 2025-01-14 09:06 | 11.2 | 4.6386 | 20.9423 | 20.4561 | 42.8107 | 4.645 | 249.4093 | 286.512 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 4.6453 | 0.14393 | 0.22791 | [8 empty] | 390.6016
CO2 | 11/07/2024 10:08:32 | 4 | 4.8697 | 21.1931 | 20.0294 | 43.5696 | 4.8835 | 248.8978 | 285.3348 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 4.8839 | 0.28407 | 0.35 | [8 empty] | 403.5089 |  | CO2 | [29 empty] | CO2 | 2025-01-14 09:08 | 11.3 | 4.6429 | 20.9588 | 20.4096 | 42.8578 | 4.6285 | 249.3587 | 286.4892 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 4.6288 | -0.30297 | 0.22191 | [8 empty] | 388.7866
CO2 | 11/07/2024 10:20:21 | 5 | 3.4118 | 21.3037 | 20.0943 | 43.8172 | 3.4177 | 248.8301 | 285.3664 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 3.418 | 0.17174 | 0.35 | 0.18205 | 0.014768 | 3 | 4.3027 | 0.036686 | 0.35 | 0.35192 | 3.41 | 280.7934 | 280.752 | CO2 | 2025-05-23 |  | 3.3334 | 21.03 | 21.58 | 42.8397 | 3.3306 | 250.0271 | 285.1136 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 3.3309 | -0.075492 | 0.27 | -0.04561 | 0.067844 | 3 | 4.3027 | 0.16853 | 0.27 | 0.31828 | 3.3333 | 279.8836 | 279.7466 | CO2 | 2025-01-14 09:24 | 13.1 | 3.542 | 22.2093 | 20.6644 | 45.7625 | 3.5442 | 248.3193 | 286.6136 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 3.5447 | 0.075493 | 0.22157 | 0.012512 | 0.080471 | 3 | 4.3027 | 0.1999 | 0.22173 | 0.29854 | 3.5427 | 278.8079 | 278.6125 | 1.47 | 8.0746 | -0.45023 | 9.8713 | 0.14039 | 11.2202 | 3.4286 | 0.039778 | 29.1661 | 0.18517 | 0.14228 | 0.29926 | 0.47542 | -0.085388 | 0.25888 | 0.32984 | -0.027266 | 0.23418 | 0.11643
CO2 | 11/07/2024 10:30:10 | 6 | 3.4099 | 21.2909 | 20.013 | 43.8056 | 3.4159 | 248.7888 | 285.3268 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 3.4162 | 0.17545 | 0.35 | [8 empty] | 280.7242 |  | CO2 | 2025-05-23 |  | 3.3328 | 21.04 | 21.67 | 42.8433 | 3.3336 | 250.074 | 285.1571 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 3.3339 | 0.032045 | 0.27 | [8 empty] | 280.1119 |  | CO2 | 2025-01-14 09:27 | 13.2 | 3.5429 | 22.202 | 20.5345 | 45.7773 | 3.5438 | 248.2448 | 286.5502 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 3.5443 | 0.040189 | 0.22219 | [8 empty] | 278.6864
CO2 | 11/07/2024 10:49:43 | 8 | 3.4082 | 21.2822 | 19.9841 | 43.7914 | 3.415 | 248.7793 | 285.3127 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 3.4153 | 0.19897 | 0.35 | [8 empty] | 280.7385 |  | CO2 | 2025-05-23 |  | 3.3336 | 21.08 | 21.69 | 42.9325 | 3.3302 | 250.049 | 285.1668 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 3.3305 | -0.093384 | 0.27 | [8 empty] | 279.2444 |  | CO2 | 2025-01-14 09:29 | 13.3 | 3.5432 | 22.192 | 20.4033 | 45.7843 | 3.5399 | 248.1702 | 286.4862 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 3.5405 | -0.078144 | 0.22143 | [8 empty] | 278.3431
CO2 | 11/07/2024 11:01:46 | 9 | 1.9557 | 21.3543 | 20.0676 | 43.9435 | 1.9583 | 248.7692 | 285.3534 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 1.9585 | 0.13101 | 0.35 | 0.14708 | 0.021924 | 3 | 4.3027 | 0.054461 | 0.35 | 0.35421 | 1.9548 | 160.4309 | 160.4213 | CO2 | 2025-05-23 |  | 1.9117 | 21 | 21.06 | 42.8837 | 1.9068 | 249.7295 | 284.862 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 1.9069 | -0.24912 | 0.27 | -0.19032 | 0.051176 | 3 | 4.3027 | 0.12713 | 0.27 | 0.29843 | 1.9099 | 160.0692 | 159.9949 | CO2 | 2025-01-14 09:37 | 14.1 | 2.1426 | 23.0235 | 21.1329 | 47.6024 | 2.1448 | 247.8442 | 286.842 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 2.1453 | 0.12602 | 0.2207 | 0.12982 | 0.0082188 | 3 | 4.3027 | 0.020417 | 0.21936 | 0.2203 | 2.1424 | 162.2072 | 162.1857 | 1.1723 | 7.9703 | -2.137 | 11.2282 | 2.6749 | 20.6043 | 2.0024 | 0.042968 | 39.8028 | 0.15851 | 0.10412 | 0.31677 | 0.32868 | -0.23329 | 0.25286 | 0.92262 | 0.086856 | 0.153 | 0.56768
CO2 | 11/07/2024 11:12:20 | 10 | 1.955 | 21.34 | 19.9692 | 43.9323 | 1.9577 | 248.7151 | 285.3055 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 1.9578 | 0.13819 | 0.35 | [8 empty] | 160.4194 |  | CO2 | 2025-05-23 |  | 1.9089 | 21 | 21.1 | 42.8749 | 1.9058 | 249.7546 | 284.8813 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 1.9059 | -0.15581 | 0.27 | [8 empty] | 160.0206 |  | CO2 | 2025-01-14 09:39 | 14.2 | 2.1425 | 23.0274 | 21.1271 | 47.6131 | 2.145 | 247.8346 | 286.8391 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 2.1455 | 0.13926 | 0.21878 | [8 empty] | 162.1846
CO2 | 11/07/2024 11:22:46 | 11 | 1.9536 | 21.336 | 19.9848 | 43.9191 | 1.957 | 248.7215 | 285.3131 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 1.9572 | 0.17206 | 0.35 | [8 empty] | 160.4136 |  | CO2 | 2025-05-23 |  | 1.9092 | 21.02 | 21.15 | 42.9108 | 1.9059 | 249.767 | 284.9055 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 1.906 | -0.16604 | 0.27 | [8 empty] | 159.8948 |  | CO2 | 2025-01-14 09:41 | 14.3 | 2.1422 | 23.023 | 21.1201 | 47.6042 | 2.1444 | 247.8377 | 286.8357 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 2.1448 | 0.1242 | 0.21859 | [8 empty] | 162.1652
CO2 | 11/07/2024 11:51:56 | 12 | 0.97692 | 21.3554 | 19.994 | 43.963 | 0.97666 | 248.7085 | 285.3176 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.97675 | -0.026954 | 0.35 | -0.054835 | 0.024597 | 3 | 4.3027 | 0.061103 | 0.35 | 0.35529 | 0.97681 | 79.9756 | 79.9646 | CO2 | 2025-05-23 |  | 0.95661 | 21.04 | 21.42 | 42.8982 | 0.95318 | 249.9175 | 285.0362 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 0.95325 | -0.3509 | 0.27 | -0.3262 | 0.027553 | 3 | 4.3027 | 0.068446 | 0.27 | 0.27854 | 0.95641 | 79.9901 | 79.9819 | CO2 | 2025-01-14 09:56 | 15.1 | 1.1444 | 23.2947 | 21.2459 | 48.2349 | 1.1459 | 247.6624 | 286.897 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 1.1461 | 0.15122 | 0.22295 | 0.13625 | 0.017428 | 3 | 4.3027 | 0.043294 | 0.22204 | 0.22623 | 1.1446 | 85.5226 | 85.5103 | -0.43439 | 7.9218 | -4.2044 | 12.8891 | 2.6622 | 19.5396 | 1.026 | -0.048983 | 40.3505 | 0.15743 | -0.0058513 | 0.31851 | 0.018371 | -0.27721 | 0.22979 | 1.2064 | 0.18523 | 0.16247 | 1.1401
CO2 | 11/07/2024 12:11:07 | 13 | 0.97685 | 21.3516 | 20.0053 | 43.9516 | 0.97622 | 248.7309 | 285.3231 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.97631 | -0.064084 | 0.35 | [8 empty] | 79.9605 |  | CO2 | 2025-05-23 |  | 0.95631 | 21.04 | 21.4 | 42.9026 | 0.9534 | 249.9049 | 285.0265 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 0.95348 | -0.29648 | 0.27 | [8 empty] | 80.0004 |  | CO2 | 2025-01-14 09:58 | 15.2 | 1.1448 | 23.3057 | 21.2844 | 48.2489 | 1.1458 | 247.6629 | 286.9158 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 1.1461 | 0.11712 | 0.22147 | [8 empty] | 85.4942
CO2 | 11/07/2024 12:30:48 | 14 | 0.97668 | 21.3482 | 20.0146 | 43.9413 | 0.97596 | 248.7309 | 285.3276 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.97605 | -0.073466 | 0.35 | [8 empty] | 79.9579 |  | CO2 | 2025-05-23 |  | 0.9563 | 21.04 | 21.36 | 42.9114 | 0.95305 | 249.8799 | 285.0072 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 0.95313 | -0.33121 | 0.27 | [8 empty] | 79.9552 |  | CO2 | 2025-01-14 10:00 | 15.3 | 1.1448 | 23.3107 | 21.3251 | 48.2493 | 1.1461 | 247.6959 | 286.9356 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 1.1464 | 0.14041 | 0.22171 | [8 empty] | 85.514
CO2 | 11/07/2024 12:51:35 | 15 | 0.49348 | 21.3527 | 19.971 | 43.9621 | 0.49293 | 248.7054 | 285.3064 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.49297 | -0.11243 | 0.35 | -0.1507 | 0.033516 | 3 | 4.3027 | 0.083258 | 0.35 | 0.35977 | 0.49355 | 40.3652 | 40.3588 | CO2 | 2025-05-23 |  | 0.4781 | 20.81 | 21.45 | 42.3541 | 0.47703 | 250.1523 | 285.0507 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0014 | 0.47706 | -0.21562 | 0.27 | -0.28629 | 0.063533 | 3 | 4.3027 | 0.15782 | 0.27 | 0.31274 | 0.47822 | 40.5467 | 40.5551 | CO2 | 2025-01-14 10:20 | 17.1 | 0.46337 | 23.0534 | 21.1376 | 47.6745 | 0.46317 | 247.8218 | 286.8442 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.46327 | -0.020693 | 0.23276 | -0.0040758 | 0.014804 | 3 | 4.3027 | 0.036776 | 0.23493 | 0.2378 | 0.46283 | 34.9753 | 34.9788 | -1.1643 | 7.7261 | -2.927 | 10.2241 | -0.072079 | 17.6845 | 0.4782 | -0.11684 | 35.6347 | 0.16752 | -0.033861 | 0.31839 | 0.10635 | -0.16945 | 0.26409 | 0.64163 | 0.11276 | 0.16877 | 0.66812
CO2 | 11/07/2024 13:10:40 | 16 | 0.49349 | 21.3491 | 19.9521 | 43.9578 | 0.49267 | 248.6927 | 285.2971 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.49272 | -0.16483 | 0.35 | [8 empty] | 40.3484 |  | CO2 | 2025-05-23 |  | 0.47833 | 20.8 | 21.47 | 42.3264 | 0.47668 | 250.1742 | 285.0604 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0014 | 0.47671 | -0.33867 | 0.27 | [8 empty] | 40.5432 |  | CO2 | 2025-01-14 10:22 | 17.2 | 0.46284 | 23.0286 | 21.0942 | 47.6237 | 0.46274 | 247.8083 | 286.8231 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.46284 | 0.00075654 | 0.23608 | [8 empty] | 34.9799
CO2 | 11/07/2024 13:29:52 | 17 | 0.49367 | 21.3507 | 19.9848 | 43.954 | 0.49281 | 248.7118 | 285.3131 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.49285 | -0.17483 | 0.35 | [8 empty] | 40.3629 |  | CO2 | 2025-05-23 |  | 0.47823 | 20.79 | 21.49 | 42.2988 | 0.47675 | 250.196 | 285.0701 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0014 | 0.47678 | -0.30457 | 0.27 | [8 empty] | 40.5752 |  | CO2 | 2025-01-14 10:24 | 17.3 | 0.46228 | 23.0022 | 21.0539 | 47.5682 | 0.46222 | 247.811 | 286.8035 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.46232 | 0.0077084 | 0.23596 | [8 empty] | 34.9812
CO2 | 11/07/2024 13:50:18 | 18 | 0.25488 | 21.3517 | 19.9413 | 43.9665 | 0.25477 | 248.6863 | 285.2919 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.25479 | -0.045245 | 0.35 | -0.052924 | 0.052241 | 3 | 4.3027 | 0.12977 | 0.35 | 0.37328 | 0.25443 | 20.8603 | 20.8179 | CO2 | 2025-05-23 |  | 0.23909 | 20.82 | 21.54 | 42.3579 | 0.23836 | 250.199 | 285.0942 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0014 | 0.23838 | -0.2989 | 0.27 | -0.2753 | 0.030086 | 3 | 4.3027 | 0.074737 | 0.27 | 0.28015 | 0.23912 | 20.2583 | 20.2338 | CO2 | 2025-01-14 10:32 | 18.1 | 0.3293 | 22.8855 | 20.8632 | 47.3302 | 0.32904 | 247.7926 | 286.7105 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.32911 | -0.060156 | 0.23311 | -0.072106 | 0.016399 | 3 | 4.3027 | 0.040737 | 0.23275 | 0.23629 | 0.32891 | 25.0272 | 25.0223 | -0.37981 | 7.1766 | -3.5076 | 12.7412 | -1.2915 | 17.9109 | 0.27415 | -0.1369 | 37.8287 | 0.16259 | 0.08398 | 0.33602 | 0.24993 | -0.13839 | 0.22815 | 0.6066 | 0.064798 | 0.17146 | 0.37793
CO2 | 11/07/2024 14:09:18 | 19 | 0.25431 | 21.3554 | 19.9298 | 43.978 | 0.25404 | 248.6703 | 285.2863 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.25406 | -0.10858 | 0.35 | [8 empty] | 20.7952 |  | CO2 | 2025-05-23 |  | 0.23913 | 20.85 | 21.62 | 42.4106 | 0.23854 | 250.2208 | 285.1329 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0014 | 0.23855 | -0.24142 | 0.27 | [8 empty] | 20.2479 |  | CO2 | 2025-01-14 10:34 | 18.2 | 0.3289 | 22.8574 | 20.8114 | 47.2759 | 0.32854 | 247.7889 | 286.6853 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.3286 | -0.090801 | 0.23358 | [8 empty] | 25.0178
CO2 | 11/07/2024 14:27:59 | 20 | 0.2541 | 21.3571 | 19.9391 | 43.9799 | 0.25408 | 248.6766 | 285.2908 | 93.52 | 44.7 | 0.02235 | 1 | 0.99861 | -0.0077 | 1.0015 | 0.25411 | -0.0049463 | 0.35 | [8 empty] | 20.7983 |  | CO2 | 2025-05-23 |  | 0.23913 | 20.89 | 21.63 | 42.5018 | 0.23843 | 250.1895 | 285.1378 | 93.52 | 44.7 | 0.02235 | 1 | 0.99862 | -0.0077 | 1.0015 | 0.23844 | -0.28557 | 0.27 | [8 empty] | 20.1953 |  | CO2 | 2025-01-14 10:36 | 18.3 | 0.32853 | 22.8321 | 20.7686 | 47.227 | 0.32825 | 247.7917 | 286.6644 | 93.52 | 44.7 | 0.02235 | 1 | 0.9986 | -0.0077 | 1.0016 | 0.32832 | -0.065359 | 0.23156 | [8 empty] | 25.0219

## Data NG FORCE 21 bar
FORCE
Cp/Cv |  |  | 1.37 | [-]
molecular mass |  |  | 28.0135 | g/mol
NG
Test Number | [25 empty] | 0.44451
Coriolis Meter
Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average
- | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h
NG | 2025-06-03 |  | 0.081218 | 21.0233 | 21.4007 | 14.6148 | 0.080261 | 435.3532 | 346.0683 | 93.52 | 44.7 | 0.02235 | 1 | 0.99955 | -0.0077 | 1.0015 | 0.080342 | -1.0778 | 0.23 | -1.0718 | 0.0085495 | 2 | 12.7062 | 0.076814 | 0.23 | 0.24249 | 0.081229 | 19.7705 | 19.7706
NG | 2025-06-03 |  | 0.08124 | 21.0301 | 21.4255 | 14.6203 | 0.080293 | 435.3743 | 346.0829 | 93.52 | 44.7 | 0.02235 | 1 | 0.99955 | -0.0077 | 1.0015 | 0.080374 | -1.0657 | 0.23 | [8 empty] | 19.7708
NG | [8 empty] | 333.2594 | 93.52 | 44.7 | 0.02235 | 1 |  | -0.0077 | 0.99984
NG | 2025-06-03 |  | 0.16247 | 21.0265 | 21.3992 | 14.622 | 0.16184 | 435.3507 | 346.0674 | 93.52 | 44.7 | 0.02235 | 1 | 0.99955 | -0.0077 | 1.0015 | 0.162 | -0.29035 | 0.23 | -0.23207 | 0.082414 | 2 | 12.7062 | 0.74046 | 0.23 | 0.77536 | 0.16248 | 39.8445 | 39.8753
NG | 2025-06-03 |  | 0.16248 | 21.0235 | 21.3509 | 14.6176 | 0.16204 | 435.3138 | 346.039 | 93.52 | 44.7 | 0.02235 | 1 | 0.99955 | -0.0077 | 1.0015 | 0.1622 | -0.1738 | 0.23 | [8 empty] | 39.9061
NG | [8 empty] | 333.2594 | 93.52 | 44.7 | 0.02235 | 1 |  | -0.0077 | 0.99984
NG | 2025-06-03 |  | 0.32521 | 21.0329 | 21.2711 | 14.6294 | 0.32509 | 435.248 | 345.9921 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 0.32542 | 0.064479 | 0.23 | 0.064113 | 0.00051727 | 2 | 12.7062 | 0.0046475 | 0.23 | 0.23005 | 0.3253 | 79.9986 | 79.9938
NG | 2025-06-03 |  | 0.32538 | 21.0398 | 21.232 | 14.6389 | 0.32526 | 435.2138 | 345.9691 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 0.32559 | 0.063747 | 0.23 | [8 empty] | 79.989
NG | [8 empty] | 333.2594 | 93.52 | 44.7 | 0.02235 | 1 |  | -0.0077 | 0.99984
NG | 2025-06-03 |  | 0.6511 | 21.0452 | 21.169 | 14.6496 | 0.65172 | 435.1638 | 345.9322 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 0.65238 | 0.19665 | 0.23 | 0.15899 | 0.053262 | 2 | 12.7062 | 0.47854 | 0.23 | 0.53094 | 0.65115 | 160.155 | 160.1108
NG | 2025-06-03 |  | 0.6512 | 21.0492 | 21.1818 | 14.649 | 0.65134 | 435.1717 | 345.9397 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 0.65199 | 0.12133 | 0.23 | [8 empty] | 160.0666
NG | [8 empty] | 333.2594 | 93.52 | 44.7 | 0.02235 | 1 |  | -0.0077 | 0.99984
NG | 2025-06-03 |  | 1.1398 | 21.0513 | 21.1549 | 14.6507 | 1.1408 | 435.148 | 345.9239 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 1.1419 | 0.18657 | 0.23 | 0.1567 | 0.042233 | 2 | 12.7062 | 0.37945 | 0.23 | 0.44372 | 1.1399 | 280.3184 | 280.21
NG | 2025-06-03 |  | 1.14 | 21.0566 | 21.1854 | 14.6556 | 1.1403 | 435.1771 | 345.9418 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0015 | 1.1414 | 0.12684 | 0.23 | [8 empty] | 280.1015
NG | [9 empty] | 93.52 | 44.7 | 0.02235 | 1 |  | -0.0077 | 0.99984
Note on I11: Chinello, Gabriele: Calculated with REFPROP v10 for 0.98 mol CH4 0.007 mol CO2 0.013 mol N2
Note on K11: Chinello, Gabriele: Calculated with REFPROP v10 for 0.98 mol CH4 0.007 mol CO2 0.013 mol N2

## Data NG FORCE 31 bar
FORCE
Cp/Cv |  |  | 1.4 | [-]
molecular mass |  |  | 28.0135 | g/mol
NG
Test Number | [25 empty] | 0.25218
Coriolis Meter
Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average
- | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h
NG | 2025-06-03 |  | 0.12202 | 31.0133 | 21.3527 | 21.9575 | 0.12092 | 432.9585 | 349.8083 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.12113 | -0.7264 | 0.23 | -0.69899 | 0.02886 | 3 | 4.3027 | 0.071693 | 0.23 | 0.24091 | 0.12198 | 19.8251 | 19.8305
NG | 2025-06-03 |  | 0.12199 | 31.0029 | 21.3651 | 21.9482 | 0.12092 | 432.9773 | 349.8157 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.12113 | -0.7017 | 0.23 | [8 empty] | 19.8336
NG | 2025-06-03 |  | 0.12194 | 30.9956 | 21.3819 | 21.9472 | 0.12091 | 432.9857 | 349.8257 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.12112 | -0.66887 | 0.23 | [8 empty] | 19.8328
NG | 2025-06-03 |  | 0.24409 | 31.0132 | 21.3169 | 21.9602 | 0.24271 | 432.9333 | 349.7871 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.24314 | -0.38847 | 0.23 | -0.31998 | 0.065384 | 3 | 4.3027 | 0.16242 | 0.23 | 0.28157 | 0.24404 | 39.7887 | 39.813
NG | 2025-06-03 |  | 0.24405 | 31.0082 | 21.3191 | 21.9602 | 0.24299 | 432.9333 | 349.7884 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.24342 | -0.25823 | 0.23 | [8 empty] | 39.8342
NG | 2025-06-03 |  | 0.24398 | 31.0019 | 21.3334 | 21.9518 | 0.24279 | 432.9438 | 349.7969 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.24322 | -0.31325 | 0.23 | [8 empty] | 39.816
NG | 2025-06-03 |  | 0.54942 | 31.0225 | 21.3109 | 21.9686 | 0.54881 | 432.9228 | 349.7835 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.54979 | 0.066402 | 0.23 | 0.029969 | 0.039459 | 3 | 4.3027 | 0.098021 | 0.23 | 0.25002 | 0.50867 | 89.9344 | 83.2326
NG | 2025-06-03 |  | 0.48831 | 31.0183 | 21.2989 | 21.9695 | 0.48739 | 432.9144 | 349.7764 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.48826 | -0.011943 | 0.23 | [8 empty] | 79.8658
NG | 2025-06-03 |  | 0.48828 | 31.0163 | 21.2998 | 21.9695 | 0.48759 | 432.9144 | 349.7769 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.48845 | 0.035449 | 0.23 | [8 empty] | 79.8976
NG | 2025-06-03 |  | 0.97581 | 31.1059 | 22.2366 | 21.9511 | 0.97426 | 433.6824 | 350.3328 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.97599 | 0.018379 | 0.23 | 0.029397 | 0.021517 | 3 | 4.3027 | 0.053452 | 0.23 | 0.23613 | 0.97573 | 159.7786 | 159.8146
NG | 2025-06-03 |  | 0.97574 | 31.102 | 22.222 | 21.9455 | 0.97416 | 433.6678 | 350.3242 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.97589 | 0.01562 | 0.23 | [8 empty] | 159.8035
NG | 2025-06-03 |  | 0.97564 | 31.1003 | 22.2382 | 21.9437 | 0.97443 | 433.6845 | 350.3338 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 0.97617 | 0.054192 | 0.23 | [8 empty] | 159.8616
NG | 2025-06-03 |  | 1.7078 | 31.109 | 22.2554 | 21.9493 | 1.7058 | 433.6992 | 350.344 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 1.7088 | 0.061788 | 0.23 | 0.096371 | 0.041701 | 3 | 4.3027 | 0.10359 | 0.23 | 0.25225 | 1.7079 | 279.7737 | 279.8862
NG | 2025-06-03 |  | 1.7079 | 31.1139 | 22.2691 | 21.9484 | 1.7073 | 433.7075 | 350.3522 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 1.7104 | 0.14268 | 0.23 | [8 empty] | 280.036
NG | 2025-06-03 |  | 1.7081 | 31.1199 | 22.2961 | 21.9531 | 1.7065 | 433.7305 | 350.3682 | 93.52 | 44.7 | 0.02235 | 1 | 0.99954 | -0.0077 | 1.0022 | 1.7096 | 0.084645 | 0.23 | [8 empty] | 279.849
Note on I11: Chinello, Gabriele: Calculated with REFPROP v10 for 0.98 mol CH4 0.007 mol CO2 0.013 mol N2
Note on K11: Chinello, Gabriele: Calculated with REFPROP v10 for 0.98 mol CH4 0.007 mol CO2 0.013 mol N2

## Data N2 NEL
NEL
Cp/Cv | 1.47 |  | [-]
molecular mass | 28.013 |  | g/mol
N2
Test Number | T03966 | [24 empty] | 0.36876
Coriolis Meter
Fluid | TestPoint Timestamp | Test Point | M_Reference | p_abs_MUT | T_MUT | rho_MUT | M_MUT | Speed of Sound | Speed of Sound (ideal gas) | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | b | Pressure correction fp | M_MUT_Corrected | dev | U_CMC(mass) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | M Ref average | Q MUT | Q MUT average
- | - | - | kg/s | bar a | °C | kg/m³ | kg/s | m/s | m/s | Hz | mm | m | - | - |  | - | kg/s | % | % | % |  |  |  | % | % | % | kg/s | m3/h | m3/h
N2 | 2024-07-09 09:02 | 4 | 4.1019 | 38.2321 | 19.9726 | 44.1977 | 4.0968 | 356.0756 | 357.6091 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 4.1054 | 0.085789 | 0.35 | 0.14135 | 0.075478 | 3 | 4.3027 | 0.1875 | 0.35 | 0.39706 | 4.102 | 333.6933 | 333.9433
N2 | 2024-07-09 09:12 | 5 | 4.1023 | 38.2309 | 20.0203 | 44.1884 | 4.1029 | 356.1077 | 357.6381 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 4.1116 | 0.22728 | 0.35 | [8 empty] | 334.2619
N2 | 2024-07-09 09:22 | 6 | 4.1018 | 38.2272 | 20.0267 | 44.1831 | 4.0977 | 356.1141 | 357.6421 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 4.1063 | 0.11098 | 0.35 | [8 empty] | 333.8748
N2 | 2024-07-09 09:34 | 7 | 2.3556 | 38.3173 | 20.0508 | 44.2835 | 2.3539 | 356.1479 | 357.6568 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 2.3589 | 0.14058 | 0.35 | 0.14122 | 0.0054393 | 3 | 4.3027 | 0.013512 | 0.35 | 0.35026 | 2.3544 | 191.3606 | 191.3603
N2 | 2024-07-09 09:50 | 8 | 2.3541 | 38.2918 | 20.0194 | 44.2591 | 2.3526 | 356.1217 | 357.6376 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 2.3576 | 0.14695 | 0.35 | [8 empty] | 191.3576
N2 | 2024-07-09 10:11 | 10 | 2.3536 | 38.2706 | 19.9596 | 44.2445 | 2.3519 | 356.0785 | 357.6011 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 2.3568 | 0.13613 | 0.35 | [8 empty] | 191.3626
N2 | 2024-07-09 10:26 | 11 | 1.1694 | 38.303 | 19.9834 | 44.2781 | 1.1672 | 356.0984 | 357.6156 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 1.1697 | 0.024783 | 0.35 | 0.035598 | 0.0093749 | 3 | 4.3027 | 0.023288 | 0.35 | 0.35077 | 1.1687 | 94.8988 | 94.8867
N2 | 2024-07-09 10:46 | 12 | 1.1686 | 38.2906 | 19.9901 | 44.2627 | 1.1666 | 356.1025 | 357.6198 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 1.1691 | 0.040583 | 0.35 | [8 empty] | 94.8853
N2 | 2024-07-09 11:05 | 13 | 1.1682 | 38.2833 | 20.0001 | 44.2525 | 1.1662 | 356.1065 | 357.6258 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 1.1687 | 0.041426 | 0.35 | [8 empty] | 94.876
N2 | 2024-07-09 11:26 | 14 | 0.5925 | 38.2842 | 19.9488 | 44.262 | 0.59075 | 356.0745 | 357.5946 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.592 | -0.084636 | 0.35 | -0.054006 | 0.029165 | 3 | 4.3027 | 0.072451 | 0.35 | 0.35742 | 0.5922 | 48.0477 | 48.0458
N2 | 2024-07-09 11:45 | 15 | 0.59216 | 38.2771 | 19.9417 | 44.255 | 0.59061 | 356.068 | 357.5902 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.59186 | -0.050813 | 0.35 | [8 empty] | 48.0442
N2 | 2024-07-09 12:04 | 16 | 0.59193 | 38.2826 | 20.0261 | 44.2474 | 0.59053 | 356.1258 | 357.6417 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.59177 | -0.026568 | 0.35 | [8 empty] | 48.0456
N2 | 2024-07-09 12:25 | 17 | 0.30288 | 38.2808 | 20.0014 | 44.2494 | 0.30136 | 356.1065 | 357.6267 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.302 | -0.2903 | 0.35 | -0.21262 | 0.067698 | 3 | 4.3027 | 0.16817 | 0.35 | 0.38831 | 0.30246 | 24.5177 | 24.5051
N2 | 2024-07-09 12:45 | 18 | 0.30233 | 38.2742 | 19.9733 | 44.2464 | 0.30114 | 356.085 | 357.6095 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.30178 | -0.18137 | 0.35 | [8 empty] | 24.5017
N2 | 2024-07-09 13:04 | 19 | 0.30218 | 38.2689 | 19.9615 | 44.2422 | 0.30104 | 356.0785 | 357.6023 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | -0.0077 | 1.0028 | 0.30168 | -0.1662 | 0.35 | [8 empty] | 24.4959

## Sheet2
0 | 1 | 1.2
11 | 1 | 1.2
Graph Mass
8.5 | 0 | 0.35968 | NEL
9 | 0 | 0.2878 | Force
9.5 | 0 | 0.27573 | DNV
Graph Volume
450 | 0 | 0 | NEL
470 | 0 | 0 | Force
490 | 0 | 0 | DNV

## Coriolis Manufact Error
From PDS
CMF 100 | Zero spec | 4.4 | kg/h | 0.0012222 | kg/s
Spec | 0.25 | From PDS
kg/s | Basic accuracy | zerro influense | Total spec | Total -Spec
0.05 | 0.25 | 2.4444 | 2.6944 | -2.6944
0.07 | 0.25 | 1.746 | 1.996 | -1.996
0.1 | 0.25 | 1.2222 | 1.4722 | -1.4722
0.15 | 0.25 | 0.81481 | 1.0648 | -1.0648
0.2 | 0.25 | 0.61111 | 0.86111 | -0.86111
0.25 | 0.25 | 0.48889 | 0.73889 | -0.73889
0.3 | 0.25 | 0.40741 | 0.65741 | -0.65741
0.35 | 0.25 | 0.34921 | 0.59921 | -0.59921
0.4 | 0.25 | 0.30556 | 0.55556 | -0.55556
0.45 | 0.25 | 0.2716 | 0.5216 | -0.5216
0.5 | 0.25 | 0.24444 | 0.49444 | -0.49444
0.55 | 0.25 | 0.22222 | 0.47222 | -0.47222
0.6 | 0.25 | 0.2037 | 0.4537 | -0.4537
0.65 | 0.25 | 0.18803 | 0.43803 | -0.43803
0.7 | 0.25 | 0.1746 | 0.4246 | -0.4246
0.75 | 0.25 | 0.16296 | 0.41296 | -0.41296
0.8 | 0.25 | 0.15278 | 0.40278 | -0.40278
0.85 | 0.25 | 0.14379 | 0.39379 | -0.39379
0.9 | 0.25 | 0.1358 | 0.3858 | -0.3858
0.95 | 0.25 | 0.12865 | 0.37865 | -0.37865
1 | 0.25 | 0.12222 | 0.37222 | -0.37222
1.05 | 0.25 | 0.1164 | 0.3664 | -0.3664
1.1 | 0.25 | 0.11111 | 0.36111 | -0.36111
1.15 | 0.25 | 0.10628 | 0.35628 | -0.35628
1.2 | 0.25 | 0.10185 | 0.35185 | -0.35185
1.25 | 0.25 | 0.097778 | 0.34778 | -0.34778
1.3 | 0.25 | 0.094017 | 0.34402 | -0.34402
1.35 | 0.25 | 0.090535 | 0.34053 | -0.34053
1.4 | 0.25 | 0.087302 | 0.3373 | -0.3373
1.45 | 0.25 | 0.084291 | 0.33429 | -0.33429
1.5 | 0.25 | 0.081481 | 0.33148 | -0.33148
1.55 | 0.25 | 0.078853 | 0.32885 | -0.32885
1.6 | 0.25 | 0.076389 | 0.32639 | -0.32639
1.65 | 0.25 | 0.074074 | 0.32407 | -0.32407
1.7 | 0.25 | 0.071895 | 0.3219 | -0.3219
1.75 | 0.25 | 0.069841 | 0.31984 | -0.31984
1.8 | 0.25 | 0.067901 | 0.3179 | -0.3179
1.85 | 0.25 | 0.066066 | 0.31607 | -0.31607
1.9 | 0.25 | 0.064327 | 0.31433 | -0.31433
1.95 | 0.25 | 0.062678 | 0.31268 | -0.31268
2 | 0.25 | 0.061111 | 0.31111 | -0.31111
2.05 | 0.25 | 0.059621 | 0.30962 | -0.30962
2.1 | 0.25 | 0.058201 | 0.3082 | -0.3082
2.15 | 0.25 | 0.056848 | 0.30685 | -0.30685
2.2 | 0.25 | 0.055556 | 0.30556 | -0.30556
2.25 | 0.25 | 0.054321 | 0.30432 | -0.30432
2.3 | 0.25 | 0.05314 | 0.30314 | -0.30314
2.35 | 0.25 | 0.052009 | 0.30201 | -0.30201
2.4 | 0.25 | 0.050926 | 0.30093 | -0.30093
2.45 | 0.25 | 0.049887 | 0.29989 | -0.29989
2.5 | 0.25 | 0.048889 | 0.29889 | -0.29889
2.55 | 0.25 | 0.04793 | 0.29793 | -0.29793
2.6 | 0.25 | 0.047009 | 0.29701 | -0.29701
2.65 | 0.25 | 0.046122 | 0.29612 | -0.29612
2.7 | 0.25 | 0.045267 | 0.29527 | -0.29527
2.75 | 0.25 | 0.044444 | 0.29444 | -0.29444
2.8 | 0.25 | 0.043651 | 0.29365 | -0.29365
2.85 | 0.25 | 0.042885 | 0.29288 | -0.29288
2.9 | 0.25 | 0.042146 | 0.29215 | -0.29215
2.95 | 0.25 | 0.041431 | 0.29143 | -0.29143
3 | 0.25 | 0.040741 | 0.29074 | -0.29074
3.05 | 0.25 | 0.040073 | 0.29007 | -0.29007
3.1 | 0.25 | 0.039427 | 0.28943 | -0.28943
3.15 | 0.25 | 0.038801 | 0.2888 | -0.2888
3.2 | 0.25 | 0.038194 | 0.28819 | -0.28819
3.25 | 0.25 | 0.037607 | 0.28761 | -0.28761
3.3 | 0.25 | 0.037037 | 0.28704 | -0.28704
3.35 | 0.25 | 0.036484 | 0.28648 | -0.28648
3.4 | 0.25 | 0.035948 | 0.28595 | -0.28595
3.45 | 0.25 | 0.035427 | 0.28543 | -0.28543
3.5 | 0.25 | 0.034921 | 0.28492 | -0.28492
3.55 | 0.25 | 0.034429 | 0.28443 | -0.28443
3.6 | 0.25 | 0.033951 | 0.28395 | -0.28395
3.65 | 0.25 | 0.033486 | 0.28349 | -0.28349
3.7 | 0.25 | 0.033033 | 0.28303 | -0.28303
3.75 | 0.25 | 0.032593 | 0.28259 | -0.28259
3.8 | 0.25 | 0.032164 | 0.28216 | -0.28216
3.85 | 0.25 | 0.031746 | 0.28175 | -0.28175
3.9 | 0.25 | 0.031339 | 0.28134 | -0.28134
3.95 | 0.25 | 0.030942 | 0.28094 | -0.28094
4 | 0.25 | 0.030556 | 0.28056 | -0.28056
4.05 | 0.25 | 0.030178 | 0.28018 | -0.28018
4.1 | 0.25 | 0.02981 | 0.27981 | -0.27981
4.15 | 0.25 | 0.029451 | 0.27945 | -0.27945
4.2 | 0.25 | 0.029101 | 0.2791 | -0.2791
4.25 | 0.25 | 0.028758 | 0.27876 | -0.27876
4.3 | 0.25 | 0.028424 | 0.27842 | -0.27842
4.35 | 0.25 | 0.028097 | 0.2781 | -0.2781
4.4 | 0.25 | 0.027778 | 0.27778 | -0.27778
4.45 | 0.25 | 0.027466 | 0.27747 | -0.27747
4.5 | 0.25 | 0.02716 | 0.27716 | -0.27716
4.55 | 0.25 | 0.026862 | 0.27686 | -0.27686
4.6 | 0.25 | 0.02657 | 0.27657 | -0.27657
4.65 | 0.25 | 0.026284 | 0.27628 | -0.27628
4.7 | 0.25 | 0.026005 | 0.276 | -0.276
4.75 | 0.25 | 0.025731 | 0.27573 | -0.27573
4.8 | 0.25 | 0.025463 | 0.27546 | -0.27546
4.85 | 0.25 | 0.0252 | 0.2752 | -0.2752
4.9 | 0.25 | 0.024943 | 0.27494 | -0.27494
4.95 | 0.25 | 0.024691 | 0.27469 | -0.27469
5 | 0.25 | 0.024444 | 0.27444 | -0.27444
5.05 | 0.25 | 0.024202 | 0.2742 | -0.2742
5.1 | 0.25 | 0.023965 | 0.27397 | -0.27397
5.15 | 0.25 | 0.023732 | 0.27373 | -0.27373
5.2 | 0.25 | 0.023504 | 0.2735 | -0.2735
5.25 | 0.25 | 0.02328 | 0.27328 | -0.27328
5.3 | 0.25 | 0.023061 | 0.27306 | -0.27306
5.35 | 0.25 | 0.022845 | 0.27285 | -0.27285
5.4 | 0.25 | 0.022634 | 0.27263 | -0.27263
5.45 | 0.25 | 0.022426 | 0.27243 | -0.27243
5.5 | 0.25 | 0.022222 | 0.27222 | -0.27222
5.55 | 0.25 | 0.022022 | 0.27202 | -0.27202
5.6 | 0.25 | 0.021825 | 0.27183 | -0.27183
5.65 | 0.25 | 0.021632 | 0.27163 | -0.27163
5.7 | 0.25 | 0.021442 | 0.27144 | -0.27144
5.75 | 0.25 | 0.021256 | 0.27126 | -0.27126
5.8 | 0.25 | 0.021073 | 0.27107 | -0.27107
5.85 | 0.25 | 0.020893 | 0.27089 | -0.27089
5.9 | 0.25 | 0.020716 | 0.27072 | -0.27072
5.95 | 0.25 | 0.020542 | 0.27054 | -0.27054
6 | 0.25 | 0.02037 | 0.27037 | -0.27037
6.05 | 0.25 | 0.020202 | 0.2702 | -0.2702
6.1 | 0.25 | 0.020036 | 0.27004 | -0.27004
6.15 | 0.25 | 0.019874 | 0.26987 | -0.26987
6.2 | 0.25 | 0.019713 | 0.26971 | -0.26971
6.25 | 0.25 | 0.019556 | 0.26956 | -0.26956
6.3 | 0.25 | 0.0194 | 0.2694 | -0.2694
6.35 | 0.25 | 0.019248 | 0.26925 | -0.26925
6.4 | 0.25 | 0.019097 | 0.2691 | -0.2691
6.45 | 0.25 | 0.018949 | 0.26895 | -0.26895
6.5 | 0.25 | 0.018803 | 0.2688 | -0.2688
6.55 | 0.25 | 0.01866 | 0.26866 | -0.26866
6.6 | 0.25 | 0.018519 | 0.26852 | -0.26852
6.65 | 0.25 | 0.018379 | 0.26838 | -0.26838
6.7 | 0.25 | 0.018242 | 0.26824 | -0.26824
6.75 | 0.25 | 0.018107 | 0.26811 | -0.26811
6.8 | 0.25 | 0.017974 | 0.26797 | -0.26797
6.85 | 0.25 | 0.017843 | 0.26784 | -0.26784
6.9 | 0.25 | 0.017713 | 0.26771 | -0.26771
6.95 | 0.25 | 0.017586 | 0.26759 | -0.26759
7 | 0.25 | 0.01746 | 0.26746 | -0.26746
7.05 | 0.25 | 0.017336 | 0.26734 | -0.26734
7.1 | 0.25 | 0.017214 | 0.26721 | -0.26721
7.15 | 0.25 | 0.017094 | 0.26709 | -0.26709
7.2 | 0.25 | 0.016975 | 0.26698 | -0.26698
7.25 | 0.25 | 0.016858 | 0.26686 | -0.26686
7.3 | 0.25 | 0.016743 | 0.26674 | -0.26674
7.35 | 0.25 | 0.016629 | 0.26663 | -0.26663
7.4 | 0.25 | 0.016517 | 0.26652 | -0.26652
7.45 | 0.25 | 0.016406 | 0.26641 | -0.26641
7.5 | 0.25 | 0.016296 | 0.2663 | -0.2663
7.55 | 0.25 | 0.016188 | 0.26619 | -0.26619
7.6 | 0.25 | 0.016082 | 0.26608 | -0.26608
7.65 | 0.25 | 0.015977 | 0.26598 | -0.26598
7.7 | 0.25 | 0.015873 | 0.26587 | -0.26587
7.75 | 0.25 | 0.015771 | 0.26577 | -0.26577
7.8 | 0.25 | 0.01567 | 0.26567 | -0.26567
7.85 | 0.25 | 0.01557 | 0.26557 | -0.26557
7.9 | 0.25 | 0.015471 | 0.26547 | -0.26547
7.95 | 0.25 | 0.015374 | 0.26537 | -0.26537
8 | 0.25 | 0.015278 | 0.26528 | -0.26528
8.05 | 0.25 | 0.015183 | 0.26518 | -0.26518
8.1 | 0.25 | 0.015089 | 0.26509 | -0.26509
8.15 | 0.25 | 0.014997 | 0.265 | -0.265
8.2 | 0.25 | 0.014905 | 0.26491 | -0.26491
8.25 | 0.25 | 0.014815 | 0.26481 | -0.26481
8.3 | 0.25 | 0.014726 | 0.26473 | -0.26473
8.35 | 0.25 | 0.014637 | 0.26464 | -0.26464
8.4 | 0.25 | 0.01455 | 0.26455 | -0.26455
8.45 | 0.25 | 0.014464 | 0.26446 | -0.26446
8.5 | 0.25 | 0.014379 | 0.26438 | -0.26438
8.55 | 0.25 | 0.014295 | 0.26429 | -0.26429
8.6 | 0.25 | 0.014212 | 0.26421 | -0.26421
8.65 | 0.25 | 0.01413 | 0.26413 | -0.26413
8.7 | 0.25 | 0.014049 | 0.26405 | -0.26405
8.75 | 0.25 | 0.013968 | 0.26397 | -0.26397
8.8 | 0.25 | 0.013889 | 0.26389 | -0.26389
8.85 | 0.25 | 0.01381 | 0.26381 | -0.26381
8.9 | 0.25 | 0.013733 | 0.26373 | -0.26373
8.95 | 0.25 | 0.013656 | 0.26366 | -0.26366
9 | 0.25 | 0.01358 | 0.26358 | -0.26358
9.05 | 0.25 | 0.013505 | 0.26351 | -0.26351
9.1 | 0.25 | 0.013431 | 0.26343 | -0.26343
9.15 | 0.25 | 0.013358 | 0.26336 | -0.26336
9.2 | 0.25 | 0.013285 | 0.26329 | -0.26329
9.25 | 0.25 | 0.013213 | 0.26321 | -0.26321
9.3 | 0.25 | 0.013142 | 0.26314 | -0.26314
9.35 | 0.25 | 0.013072 | 0.26307 | -0.26307
9.4 | 0.25 | 0.013002 | 0.263 | -0.263
9.45 | 0.25 | 0.012934 | 0.26293 | -0.26293
9.5 | 0.25 | 0.012865 | 0.26287 | -0.26287
9.55 | 0.25 | 0.012798 | 0.2628 | -0.2628
9.6 | 0.25 | 0.012731 | 0.26273 | -0.26273
9.65 | 0.25 | 0.012666 | 0.26267 | -0.26267
9.7 | 0.25 | 0.0126 | 0.2626 | -0.2626
9.75 | 0.25 | 0.012536 | 0.26254 | -0.26254
9.8 | 0.25 | 0.012472 | 0.26247 | -0.26247
9.85 | 0.25 | 0.012408 | 0.26241 | -0.26241
9.9 | 0.25 | 0.012346 | 0.26235 | -0.26235
9.95 | 0.25 | 0.012284 | 0.26228 | -0.26228
10 | 0.25 | 0.012222 | 0.26222 | -0.26222
10.05 | 0.25 | 0.012161 | 0.26216 | -0.26216
10.1 | 0.25 | 0.012101 | 0.2621 | -0.2621
10.15 | 0.25 | 0.012042 | 0.26204 | -0.26204
10.2 | 0.25 | 0.011983 | 0.26198 | -0.26198
10.25 | 0.25 | 0.011924 | 0.26192 | -0.26192
10.3 | 0.25 | 0.011866 | 0.26187 | -0.26187
10.35 | 0.25 | 0.011809 | 0.26181 | -0.26181
10.4 | 0.25 | 0.011752 | 0.26175 | -0.26175
10.45 | 0.25 | 0.011696 | 0.2617 | -0.2617
10.5 | 0.25 | 0.01164 | 0.26164 | -0.26164
10.55 | 0.25 | 0.011585 | 0.26159 | -0.26159
10.6 | 0.25 | 0.01153 | 0.26153 | -0.26153
10.65 | 0.25 | 0.011476 | 0.26148 | -0.26148
10.7 | 0.25 | 0.011423 | 0.26142 | -0.26142
10.75 | 0.25 | 0.01137 | 0.26137 | -0.26137
10.8 | 0.25 | 0.011317 | 0.26132 | -0.26132
10.85 | 0.25 | 0.011265 | 0.26126 | -0.26126
10.9 | 0.25 | 0.011213 | 0.26121 | -0.26121
10.95 | 0.25 | 0.011162 | 0.26116 | -0.26116
11 | 0.25 | 0.011111 | 0.26111 | -0.26111
11.05 | 0.25 | 0.011061 | 0.26106 | -0.26106
11.1 | 0.25 | 0.011011 | 0.26101 | -0.26101
11.15 | 0.25 | 0.010962 | 0.26096 | -0.26096
11.2 | 0.25 | 0.010913 | 0.26091 | -0.26091
11.25 | 0.25 | 0.010864 | 0.26086 | -0.26086
11.3 | 0.25 | 0.010816 | 0.26082 | -0.26082
11.35 | 0.25 | 0.010768 | 0.26077 | -0.26077
11.4 | 0.25 | 0.010721 | 0.26072 | -0.26072
11.45 | 0.25 | 0.010674 | 0.26067 | -0.26067
11.5 | 0.25 | 0.010628 | 0.26063 | -0.26063
11.55 | 0.25 | 0.010582 | 0.26058 | -0.26058
11.6 | 0.25 | 0.010536 | 0.26054 | -0.26054
11.65 | 0.25 | 0.010491 | 0.26049 | -0.26049
11.7 | 0.25 | 0.010446 | 0.26045 | -0.26045
11.75 | 0.25 | 0.010402 | 0.2604 | -0.2604
11.8 | 0.25 | 0.010358 | 0.26036 | -0.26036
11.85 | 0.25 | 0.010314 | 0.26031 | -0.26031
11.9 | 0.25 | 0.010271 | 0.26027 | -0.26027
11.95 | 0.25 | 0.010228 | 0.26023 | -0.26023
12 | 0.25 | 0.010185 | 0.26019 | -0.26019
12.05 | 0.25 | 0.010143 | 0.26014 | -0.26014
12.1 | 0.25 | 0.010101 | 0.2601 | -0.2601
12.15 | 0.25 | 0.010059 | 0.26006 | -0.26006
12.2 | 0.25 | 0.010018 | 0.26002 | -0.26002
12.25 | 0.25 | 0.0099773 | 0.25998 | -0.25998
12.3 | 0.25 | 0.0099368 | 0.25994 | -0.25994
12.35 | 0.25 | 0.0098965 | 0.2599 | -0.2599
12.4 | 0.25 | 0.0098566 | 0.25986 | -0.25986
12.45 | 0.25 | 0.009817 | 0.25982 | -0.25982
12.5 | 0.25 | 0.0097778 | 0.25978 | -0.25978
12.55 | 0.25 | 0.0097388 | 0.25974 | -0.25974
12.6 | 0.25 | 0.0097002 | 0.2597 | -0.2597
12.65 | 0.25 | 0.0096618 | 0.25966 | -0.25966
12.7 | 0.25 | 0.0096238 | 0.25962 | -0.25962
12.75 | 0.25 | 0.0095861 | 0.25959 | -0.25959
12.8 | 0.25 | 0.0095486 | 0.25955 | -0.25955
12.85 | 0.25 | 0.0095115 | 0.25951 | -0.25951
12.9 | 0.25 | 0.0094746 | 0.25947 | -0.25947
12.95 | 0.25 | 0.009438 | 0.25944 | -0.25944
13 | 0.25 | 0.0094017 | 0.2594 | -0.2594
13.05 | 0.25 | 0.0093657 | 0.25937 | -0.25937
13.1 | 0.25 | 0.0093299 | 0.25933 | -0.25933
13.15 | 0.25 | 0.0092945 | 0.25929 | -0.25929
13.2 | 0.25 | 0.0092593 | 0.25926 | -0.25926
13.25 | 0.25 | 0.0092243 | 0.25922 | -0.25922
13.3 | 0.25 | 0.0091896 | 0.25919 | -0.25919
13.35 | 0.25 | 0.0091552 | 0.25916 | -0.25916
13.4 | 0.25 | 0.0091211 | 0.25912 | -0.25912
13.45 | 0.25 | 0.0090872 | 0.25909 | -0.25909
13.5 | 0.25 | 0.0090535 | 0.25905 | -0.25905
13.55 | 0.25 | 0.0090201 | 0.25902 | -0.25902
13.6 | 0.25 | 0.0089869 | 0.25899 | -0.25899
13.65 | 0.25 | 0.008954 | 0.25895 | -0.25895
13.7 | 0.25 | 0.0089213 | 0.25892 | -0.25892
13.75 | 0.25 | 0.0088889 | 0.25889 | -0.25889
13.8 | 0.25 | 0.0088567 | 0.25886 | -0.25886
13.85 | 0.25 | 0.0088247 | 0.25882 | -0.25882
13.9 | 0.25 | 0.008793 | 0.25879 | -0.25879
13.95 | 0.25 | 0.0087614 | 0.25876 | -0.25876
14 | 0.25 | 0.0087302 | 0.25873 | -0.25873
14.05 | 0.25 | 0.0086991 | 0.2587 | -0.2587
14.1 | 0.25 | 0.0086682 | 0.25867 | -0.25867
14.15 | 0.25 | 0.0086376 | 0.25864 | -0.25864
14.2 | 0.25 | 0.0086072 | 0.25861 | -0.25861
14.25 | 0.25 | 0.008577 | 0.25858 | -0.25858
14.3 | 0.25 | 0.008547 | 0.25855 | -0.25855
14.35 | 0.25 | 0.0085172 | 0.25852 | -0.25852
14.4 | 0.25 | 0.0084877 | 0.25849 | -0.25849
14.45 | 0.25 | 0.0084583 | 0.25846 | -0.25846
14.5 | 0.25 | 0.0084291 | 0.25843 | -0.25843
14.55 | 0.25 | 0.0084002 | 0.2584 | -0.2584
14.6 | 0.25 | 0.0083714 | 0.25837 | -0.25837
14.65 | 0.25 | 0.0083428 | 0.25834 | -0.25834
14.7 | 0.25 | 0.0083144 | 0.25831 | -0.25831
14.75 | 0.25 | 0.0082863 | 0.25829 | -0.25829
14.8 | 0.25 | 0.0082583 | 0.25826 | -0.25826
14.85 | 0.25 | 0.0082305 | 0.25823 | -0.25823
14.9 | 0.25 | 0.0082028 | 0.2582 | -0.2582
14.95 | 0.25 | 0.0081754 | 0.25818 | -0.25818
15 | 0.25 | 0.0081481 | 0.25815 | -0.25815
15.05 | 0.25 | 0.0081211 | 0.25812 | -0.25812
15.1 | 0.25 | 0.0080942 | 0.25809 | -0.25809
15.15 | 0.25 | 0.0080675 | 0.25807 | -0.25807
15.2 | 0.25 | 0.0080409 | 0.25804 | -0.25804
15.25 | 0.25 | 0.0080146 | 0.25801 | -0.25801
15.3 | 0.25 | 0.0079884 | 0.25799 | -0.25799
15.35 | 0.25 | 0.0079624 | 0.25796 | -0.25796
15.4 | 0.25 | 0.0079365 | 0.25794 | -0.25794
15.45 | 0.25 | 0.0079108 | 0.25791 | -0.25791
15.5 | 0.25 | 0.0078853 | 0.25789 | -0.25789
15.55 | 0.25 | 0.0078599 | 0.25786 | -0.25786
15.6 | 0.25 | 0.0078348 | 0.25783 | -0.25783
15.65 | 0.25 | 0.0078097 | 0.25781 | -0.25781
15.7 | 0.25 | 0.0077849 | 0.25778 | -0.25778
15.75 | 0.25 | 0.0077601 | 0.25776 | -0.25776
15.8 | 0.25 | 0.0077356 | 0.25774 | -0.25774
15.85 | 0.25 | 0.0077112 | 0.25771 | -0.25771
15.9 | 0.25 | 0.0076869 | 0.25769 | -0.25769
15.95 | 0.25 | 0.0076628 | 0.25766 | -0.25766
16 | 0.25 | 0.0076389 | 0.25764 | -0.25764
16.05 | 0.25 | 0.0076151 | 0.25762 | -0.25762
16.1 | 0.25 | 0.0075914 | 0.25759 | -0.25759
16.15 | 0.25 | 0.0075679 | 0.25757 | -0.25757
16.2 | 0.25 | 0.0075446 | 0.25754 | -0.25754
16.25 | 0.25 | 0.0075214 | 0.25752 | -0.25752
16.3 | 0.25 | 0.0074983 | 0.2575 | -0.2575
16.35 | 0.25 | 0.0074754 | 0.25748 | -0.25748
16.4 | 0.25 | 0.0074526 | 0.25745 | -0.25745
16.45 | 0.25 | 0.0074299 | 0.25743 | -0.25743
16.5 | 0.25 | 0.0074074 | 0.25741 | -0.25741
16.55 | 0.25 | 0.007385 | 0.25739 | -0.25739
16.6 | 0.25 | 0.0073628 | 0.25736 | -0.25736
16.65 | 0.25 | 0.0073407 | 0.25734 | -0.25734
16.7 | 0.25 | 0.0073187 | 0.25732 | -0.25732
16.75 | 0.25 | 0.0072968 | 0.2573 | -0.2573
16.8 | 0.25 | 0.0072751 | 0.25728 | -0.25728
16.85 | 0.25 | 0.0072535 | 0.25725 | -0.25725
16.9 | 0.25 | 0.0072321 | 0.25723 | -0.25723
16.95 | 0.25 | 0.0072108 | 0.25721 | -0.25721
17 | 0.25 | 0.0071895 | 0.25719 | -0.25719
17.05 | 0.25 | 0.0071685 | 0.25717 | -0.25717
17.1 | 0.25 | 0.0071475 | 0.25715 | -0.25715
17.15 | 0.25 | 0.0071267 | 0.25713 | -0.25713
17.2 | 0.25 | 0.0071059 | 0.25711 | -0.25711
17.25 | 0.25 | 0.0070853 | 0.25709 | -0.25709
17.3 | 0.25 | 0.0070649 | 0.25706 | -0.25706
17.35 | 0.25 | 0.0070445 | 0.25704 | -0.25704
17.4 | 0.25 | 0.0070243 | 0.25702 | -0.25702
17.45 | 0.25 | 0.0070041 | 0.257 | -0.257
17.5 | 0.25 | 0.0069841 | 0.25698 | -0.25698
17.55 | 0.25 | 0.0069642 | 0.25696 | -0.25696
17.6 | 0.25 | 0.0069444 | 0.25694 | -0.25694

## (TBL) Coriolis N2 35 bar
Table: Coriolis | [57 empty] | a0 | 0.08203
[58 empty] | a1 | -0.012922
Job Number | EMR013 | [49 empty] | 4239750.0944 | [6 empty] | a2 | -0.11908
Test Number | T03127 | [49 empty] | 108050.235 | [6 empty] | a3 | -0.067147
Test Title | 3" Coriolis Intercomparison 35 Bar | [25 empty] | 0 | [30 empty] | a4 | 0.1215
Test Description | 3" Coriolis Intercomparison 35 Bar | [56 empty] | b0 | -0.00011761
[39 empty] | 0.21677 |  |  |  | 0.25858 | [14 empty] | b1 | -0.016601
[43 empty] | 0.44189 | [14 empty] | cp | -0.042815
ONLY PRESSURE CORRECTION APPLIED - NO COMPRESSIBILITY
[15 empty] | RefProp version 10 | [13 empty] | Averaged values | [8 empty] | Averaged values | [5 empty] | Reference Turbine Calcs
| Test Point | TestPoint Timestamp | Collection Time (t) | Ambient Pressure (Pamb) | Test Temperature (T) | Test Pressure (Pgauge) | Reference Mass Flow Rate (ṁ) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | K Factor (K) | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Test Temperature (T) | Test Pressure (Pabs) | Gas Density (ρ) | Speed of Sound | Isentropic expansion | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | Pressure correction fp | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Test Temperature (T) | Test Pressure (Pgauge) | Meter Est Volume Flow (Qi) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | Reynolds | K nominal | Meter Nominal Flow (Q) | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Meter Est Volume Flow (Qi) | Reynolds | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Reference Volume Flow Rate (Q) | Reference Mass Flow Rate (ṁ)
| - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | % | °C | bara | kg/m³ | m/s | - | Hz | mm | mm | - | - | - | kg/s | % | kg/s | kg/s | % | % | % | % | % | kg/s | % | kg/s | kg/s | % | % | % | % | % | °C | barg | m³/h | kg/m³ | Pa s | pulse | Hz | - | - | m3/s | - | % | % | % | % | pulse/m³ | m³/h | - | - | % | % | % | % | pulse/m³ | m³/h | kg/s
| 2 | 2023-11-08 07:34 | 181.7437 | 0.96856 | 20.1175 | 35.7537 | 9.2206 | 42.4257 | 1.8185e-05 | 177094 | 974.4163 | 105.6958 | 9.7442 | 5.6785 | 20.1175 | 36.7223 | 42.4267 | 355.8208 | 1.4628 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 9.7629 | 5.8812 | 9.2206 |  |  | 1.7984 |  |  |  | 9.7695 | 5.9533 | 9.2206 |  |  | 1.7996 |  |  |  | 20.1808 | 34.6641 | 806.4454 | 41.1542 | 1.8167e-05 | 236085 | 1298.9998 | 4239750.0944 | 5798.78 | 806.4454 | 0.62734 | 0.029298 | -0.0018576 | -0.011077 | 0.016363 | 5797.8311 | 806.5773 | 4240443.9601 | 0.62741 | 0.029289 | -0.0018573 | -0.011079 | 0.016353 | 5797.8317 | 806.5773 | 9.2206
| 3 | 2023-11-08 07:37 | 184.4997 | 0.96852 | 20.0799 | 35.7603 | 9.2208 | 42.4391 | 1.8183e-05 | 182952 | 991.6114 | 107.5579 | 9.9161 | 7.5403 | 20.0799 | 36.7288 | 42.4401 | 355.7982 | 1.4628 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 9.9351 | 7.7466 | 9.2208 | [6 empty] | 9.9419 | 7.82 | 9.2208 | [6 empty] | 20.2319 | 34.6703 | 806.4821 | 41.1535 | 1.8169e-05 | 239676 | 1299.059 | 4239319.9025 | 5798.78 | 806.4821 | 0.6273 | 0.029303 | -0.0018576 | -0.011076 | 0.01637 | 5797.8307 | 806.6142 | 4240013.9949 | 0.62737 | 0.029295 | -0.0018573 | -0.011078 | 0.01636 | 5797.8313 | 806.6141 | 9.2208
| 4 | 2023-11-08 07:41 | 181.7772 | 0.96849 | 20.0829 | 35.7669 | 9.2222 | 42.4463 | 1.8184e-05 | 174262 | 958.657 | 103.9682 | 9.5866 | 3.9512 | 20.0829 | 36.7354 | 42.4473 | 355.8017 | 1.4628 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 9.605 | 4.1507 | 9.2222 | [6 empty] | 9.6115 | 4.2216 | 9.2222 | [6 empty] | 20.264 | 34.6769 | 806.5459 | 41.1563 | 1.817e-05 | 236158 | 1299.1617 | 4239582.3248 | 5798.78 | 806.5459 | 0.62732 | 0.0293 | -0.0018573 | -0.011075 | 0.016367 | 5797.8309 | 806.6779 | 4240276.3517 | 0.62739 | 0.029292 | -0.001857 | -0.011077 | 0.016357 | 5797.8315 | 806.6778 | 9.2222
| 6 | 2023-11-08 07:53 | 186.1061 | 0.96847 | 20.0769 | 35.7139 | 8.1783 | 42.3858 | 1.8182e-05 | 152384 | 818.8016 | 100.1431 | 8.188 | 0.11925 | 20.0769 | 36.6823 | 42.3868 | 355.7856 | 1.4627 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 8.2037 | 0.31096 | 8.1783 |  |  | 0.20841 |  |  |  | 8.2093 | 0.3793 | 8.1783 |  |  | 0.20856 |  |  |  | 20.1982 | 34.8653 | 711.2511 | 41.3844 | 1.8171e-05 | 213215 | 1145.6635 | 3759240.1575 | 5798.78 | 711.2511 | 0.5751 | 0.035732 | -0.0021031 | -0.0097693 | 0.023859 | 5797.3964 | 711.4208 | 3760137.3014 | 0.5752 | 0.035719 | -0.0021026 | -0.0097717 | 0.023845 | 5797.3973 | 711.4207 | 8.1783
| 7 | 2023-11-08 07:57 | 186.8649 | 0.9686 | 20.0028 | 35.7021 | 8.1766 | 42.384 | 1.8179e-05 | 153514 | 821.5238 | 100.4967 | 8.2152 | 0.47269 | 20.0028 | 36.6707 | 42.385 | 355.7355 | 1.4627 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 8.231 | 0.66497 | 8.1766 | [6 empty] | 8.2366 | 0.73357 | 8.1766 | [6 empty] | 20.1722 | 34.8535 | 711.2699 | 41.3749 | 1.817e-05 | 214090 | 1145.6937 | 3758752.115 | 5798.78 | 711.2699 | 0.57504 | 0.035739 | -0.0021036 | -0.0097705 | 0.023865 | 5797.3961 | 711.4396 | 3759649.3446 | 0.57515 | 0.035726 | -0.0021031 | -0.0097729 | 0.02385 | 5797.397 | 711.4395 | 8.1766
| 8 | 2023-11-08 08:00 | 186.6627 | 0.96877 | 19.9956 | 35.6977 | 8.1767 | 42.3803 | 1.8178e-05 | 152791 | 818.5404 | 100.1297 | 8.1854 | 0.10585 | 19.9956 | 36.6665 | 42.3813 | 355.7299 | 1.4627 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 8.2011 | 0.29739 | 8.1767 | [6 empty] | 8.2067 | 0.36574 | 8.1767 | [6 empty] | 20.1459 | 34.8493 | 711.2953 | 41.3742 | 1.8169e-05 | 213866 | 1145.7348 | 3759081.2302 | 5798.78 | 711.2953 | 0.57508 | 0.035734 | -0.0021035 | -0.0097718 | 0.023859 | 5797.3965 | 711.4651 | 3759978.3136 | 0.57519 | 0.035721 | -0.002103 | -0.0097742 | 0.023844 | 5797.3973 | 711.465 | 8.1767
| 10 | 2023-11-08 08:12 | 187.2494 | 0.96934 | 19.9553 | 35.6112 | 6.5501 | 42.2871 | 1.8175e-05 | 122781 | 655.7083 | 100.1435 | 6.5571 | 0.10631 | 19.9553 | 36.5805 | 42.288 | 355.6844 | 1.4625 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 6.5696 | 0.29717 | 6.5501 | 6.5486 | 0.26358 | 0.081736 | 0.2487 | 0.25 | 0.35263 | 6.5741 | 0.36554 | 6.5501 | 6.5486 | 0.33193 | 0.081792 | 0.24887 | 0.25 | 0.35275 | 20.0203 | 35.0728 | 565.9047 | 41.6531 | 1.8167e-05 | 170686 | 911.5436 | 3011071.2139 | 5798.78 | 565.9047 | 0.47872 | 0.047568 | -0.0026496 | -0.0077782 | 0.03714 | 5796.6263 | 566.115 | 3012189.9434 | 0.47888 | 0.047549 | -0.0026486 | -0.0077811 | 0.037119 | 5796.6276 | 566.1148 | 6.5501
| 11 | 2023-11-08 08:16 | 187.4395 | 0.9695 | 19.9354 | 35.5998 | 6.548 | 42.2772 | 1.8174e-05 | 122710 | 654.6644 | 100.017 | 6.5466 | -0.020125 | 19.9354 | 36.5693 | 42.2782 | 355.6691 | 1.4625 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 6.5591 | 0.1704 | 6.548 | [6 empty] | 6.5636 | 0.23868 | 6.548 | [6 empty] | 20.0009 | 35.0616 | 565.8507 | 41.6433 | 1.8166e-05 | 170843 | 911.4565 | 3010255.0488 | 5798.78 | 565.8507 | 0.4786 | 0.047582 | -0.0026505 | -0.007778 | 0.037153 | 5796.6256 | 566.061 | 3011373.8789 | 0.47876 | 0.047563 | -0.0026495 | -0.0077809 | 0.037132 | 5796.6268 | 566.0608 | 6.548
| 12 | 2023-11-08 08:19 | 187.552 | 0.96934 | 19.9211 | 35.5918 | 6.5478 | 42.27 | 1.8173e-05 | 122967 | 655.6421 | 100.1696 | 6.5564 | 0.13242 | 19.9211 | 36.5612 | 42.271 | 355.658 | 1.4625 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 6.5689 | 0.32316 | 6.5478 | [6 empty] | 6.5734 | 0.39156 | 6.5478 | [6 empty] | 19.9835 | 35.0535 | 565.9251 | 41.6365 | 1.8165e-05 | 170968 | 911.5765 | 3010311.1986 | 5798.78 | 565.9251 | 0.47861 | 0.047581 | -0.0026506 | -0.0077795 | 0.037151 | 5796.6257 | 566.1355 | 3011429.9723 | 0.47877 | 0.047562 | -0.0026495 | -0.0077824 | 0.03713 | 5796.6269 | 566.1353 | 6.5478
| 14 | 2023-11-08 08:35 | 187.5932 | 0.96965 | 19.8922 | 35.5103 | 4.8134 | 42.1804 | 1.817e-05 | 90205 | 480.8544 | 99.9523 | 4.8085 | -0.10065 | 19.8922 | 36.48 | 42.1814 | 355.6209 | 1.4623 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 4.8177 | 0.089002 | 4.8134 | 4.8126 | 0.11369 | 0.058263 | 0.17727 | 0.25 | 0.30647 | 4.821 | 0.15725 | 4.8134 | 4.8126 | 0.18196 | 0.058303 | 0.1774 | 0.25 | 0.30655 | 19.8771 | 35.2257 | 413.807 | 41.8529 | 1.8164e-05 | 125040 | 666.5488 | 2212753.4495 | 5798.78 | 413.807 | 0.34493 | 0.062369 | -0.0036634 | -0.0056907 | 0.053015 | 5795.7058 | 414.0265 | 2213927.1585 | 0.34516 | 0.062346 | -0.0036613 | -0.0056938 | 0.052991 | 5795.7072 | 414.0264 | 4.8134
| 15 | 2023-11-08 08:39 | 181.8656 | 0.96972 | 19.904 | 35.5035 | 4.8124 | 42.1708 | 1.8171e-05 | 87418 | 480.6737 | 99.9352 | 4.8067 | -0.11774 | 19.904 | 36.4733 | 42.1718 | 355.627 | 1.4623 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 4.8159 | 0.071838 | 4.8124 | [6 empty] | 4.8191 | 0.14007 | 4.8124 | [6 empty] | 19.8755 | 35.219 | 413.7958 | 41.8454 | 1.8164e-05 | 121219 | 666.5307 | 2212324.055 | 5798.78 | 413.7958 | 0.34485 | 0.062377 | -0.0036641 | -0.0056906 | 0.053023 | 5795.7053 | 414.0153 | 2213497.7072 | 0.34508 | 0.062354 | -0.0036621 | -0.0056936 | 0.052999 | 5795.7067 | 414.0152 | 4.8124
| 16 | 2023-11-08 08:42 | 181.7865 | 0.96967 | 19.8972 | 35.4965 | 4.8121 | 42.1636 | 1.817e-05 | 87470 | 481.1689 | 100.0435 | 4.8117 | -0.0094902 | 19.8972 | 36.4662 | 42.1646 | 355.621 | 1.4623 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 4.8208 | 0.18023 | 4.8121 | [6 empty] | 4.8241 | 0.24855 | 4.8121 | [6 empty] | 19.8728 | 35.2121 | 413.8495 | 41.8378 | 1.8164e-05 | 121182 | 666.6173 | 2212236.9695 | 5798.78 | 413.8495 | 0.34483 | 0.062379 | -0.0036643 | -0.0056914 | 0.053023 | 5795.7053 | 414.0691 | 2213410.5918 | 0.34506 | 0.062356 | -0.0036622 | -0.0056944 | 0.052999 | 5795.7067 | 414.069 | 4.8121
| 18 | 2023-11-08 08:56 | 182.3775 | 0.96985 | 19.838 | 35.4398 | 3.297 | 42.1074 | 1.8166e-05 | 60034 | 329.1743 | 99.9065 | 3.2917 | -0.15963 | 19.838 | 36.4096 | 42.1084 | 355.5702 | 1.4622 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 3.298 | 0.029347 | 3.297 | 3.2964 | -0.0034834 | 0.060202 | 0.18318 | 0.25 | 0.30993 | 3.3002 | 0.097576 | 3.297 | 3.2964 | 0.064715 | 0.060246 | 0.18331 | 0.25 | 0.31 | 19.7747 | 35.3126 | 282.6169 | 41.9698 | 1.8161e-05 | 83024 | 455.2315 | 1515694.5656 | 5798.78 | 282.6169 | 0.18061 | 0.075546 | -0.0054931 | -0.0038881 | 0.066164 | 5794.9433 | 282.8041 | 1516698.0802 | 0.1809 | 0.075528 | -0.0054892 | -0.0038907 | 0.066149 | 5794.9442 | 282.804 | 3.297
| 19 | 2023-11-08 08:59 | 182.3203 | 0.96997 | 19.8874 | 35.4387 | 3.2963 | 42.0985 | 1.8169e-05 | 59941 | 328.7676 | 99.8043 | 3.2877 | -0.26176 | 19.8874 | 36.4087 | 42.0995 | 355.6016 | 1.4622 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 3.2939 | -0.072964 | 3.2963 | [6 empty] | 3.2961 | -0.0048168 | 3.2963 | [6 empty] | 19.7773 | 35.3116 | 282.5661 | 41.9684 | 1.8161e-05 | 82983 | 455.1496 | 1515363.5765 | 5798.78 | 282.5661 | 0.18052 | 0.075551 | -0.0054944 | -0.0038874 | 0.06617 | 5794.943 | 282.7532 | 1516366.95 | 0.1808 | 0.075534 | -0.0054904 | -0.0038899 | 0.066154 | 5794.9439 | 282.7531 | 3.2963
| 20 | 2023-11-08 09:03 | 182.8671 | 0.96996 | 19.895 | 35.4338 | 3.296 | 42.0917 | 1.8169e-05 | 60179 | 329.0861 | 99.9103 | 3.2909 | -0.15579 | 19.895 | 36.4038 | 42.0926 | 355.6053 | 1.4622 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 3.2971 | 0.033167 | 3.296 | [6 empty] | 3.2993 | 0.10139 | 3.296 | [6 empty] | 19.7876 | 35.3065 | 282.5903 | 41.9609 | 1.8162e-05 | 83239 | 455.1886 | 1515192.4126 | 5798.78 | 282.5903 | 0.18047 | 0.075554 | -0.0054948 | -0.0038875 | 0.066172 | 5794.9428 | 282.7774 | 1516195.7071 | 0.18076 | 0.075537 | -0.0054909 | -0.0038901 | 0.066156 | 5794.9438 | 282.7774 | 3.296
| 22 | 2023-11-08 09:14 | 182.7371 | 0.96997 | 19.7447 | 35.3888 | 1.5997 | 42.0631 | 1.8161e-05 | 29209 | 159.8417 | 99.9849 | 1.5984 | -0.08272 | 19.7447 | 36.3588 | 42.0641 | 355.4988 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 1.6014 | 0.10599 | 1.5997 | 1.5999 | 0.078762 | 0.036179 | 0.11008 | 0.25 | 0.27316 | 1.6025 | 0.17429 | 1.5997 | 1.5999 | 0.14704 | 0.036217 | 0.1102 | 0.25 | 0.27321 | 19.6133 | 35.3696 | 136.8279 | 42.0614 | 1.8155e-05 | 40275 | 220.3986 | 735660.7934 | 5798.78 | 136.8279 | -0.13332 | 0.081834 | -0.01232 | -0.0018835 | 0.067631 | 5794.8582 | 136.9205 | 736158.6633 | -0.13303 | 0.081838 | -0.01231 | -0.0018848 | 0.067643 | 5794.8575 | 136.9205 | 1.5997
| 23 | 2023-11-08 09:18 | 183.8976 | 0.97005 | 19.8545 | 35.3946 | 1.6001 | 42.0526 | 1.8166e-05 | 29381 | 159.7683 | 99.9167 | 1.5977 | -0.15094 | 19.8545 | 36.3646 | 42.0536 | 355.5704 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 1.6007 | 0.037708 | 1.6001 | [6 empty] | 1.6018 | 0.10594 | 1.6001 | [6 empty] | 19.6016 | 35.375 | 136.8321 | 42.0695 | 1.8155e-05 | 40532 | 220.4053 | 735841.8947 | 5798.78 | 136.8321 | -0.13322 | 0.081836 | -0.012317 | -0.0018837 | 0.067635 | 5794.858 | 136.9247 | 736339.918 | -0.13292 | 0.08184 | -0.012307 | -0.001885 | 0.067647 | 5794.8573 | 136.9247 | 1.6001
| 24 | 2023-11-08 09:21 | 182.9094 | 0.97008 | 19.7978 | 35.3826 | 1.5998 | 42.0477 | 1.8164e-05 | 29234 | 159.8278 | 99.9716 | 1.5983 | -0.096052 | 19.7978 | 36.3527 | 42.0486 | 355.5314 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 1.6013 | 0.092592 | 1.5998 | [6 empty] | 1.6024 | 0.16088 | 1.5998 | [6 empty] | 19.5996 | 35.363 | 136.8518 | 42.056 | 1.8155e-05 | 40320 | 220.4371 | 735724.5082 | 5798.78 | 136.8518 | -0.13328 | 0.081835 | -0.012319 | -0.001884 | 0.067632 | 5794.8582 | 136.9444 | 736222.4288 | -0.13299 | 0.081839 | -0.012309 | -0.0018852 | 0.067644 | 5794.8575 | 136.9444 | 1.5998
| 26 | 2023-11-08 09:29 | 182.9186 | 0.97021 | 19.8099 | 35.3844 | 0.5923 | 42.048 | 1.8164e-05 | 10831 | 59.2121 | 100.0024 | 0.59212 | -0.030586 | 19.8099 | 36.3546 | 42.049 | 355.5396 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.59324 | 0.1582 | 0.5923 | 0.59238 | 0.16842 | 0.02169 | 0.065997 | 0.25 | 0.25856 | 0.59364 | 0.22653 | 0.5923 | 0.59238 | 0.23676 | 0.02171 | 0.066057 | 0.25 | 0.25858 | 19.5752 | 35.3953 | 50.6346 | 42.0974 | 1.8154e-05 | 14919 | 81.5609 | 272489.8587 | 5798.78 | 50.6346 | -0.56465 | 0.075799 | -0.042159 | -0.00069713 | 0.032943 | 5796.8697 | 50.6513 | 272579.6541 | -0.56451 | 0.075794 | -0.04214 | -0.00069736 | 0.032957 | 5796.8689 | 50.6513 | 0.5923
| 27 | 2023-11-08 09:33 | 183.071 | 0.97023 | 19.8099 | 35.385 | 0.5926 | 42.0487 | 1.8164e-05 | 10845 | 59.2393 | 99.9979 | 0.59239 | -0.035053 | 19.8099 | 36.3552 | 42.0497 | 355.5397 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.59351 | 0.15373 | 0.5926 | [6 empty] | 0.59392 | 0.22205 | 0.5926 | [6 empty] | 19.5423 | 35.3955 | 50.6535 | 42.1029 | 1.8153e-05 | 14937 | 81.5913 | 272648.6783 | 5798.78 | 50.6535 | -0.5644 | 0.075791 | -0.042132 | -0.00069747 | 0.032961 | 5796.8686 | 50.6702 | 272738.5765 | -0.56425 | 0.075787 | -0.042114 | -0.0006977 | 0.032975 | 5796.8678 | 50.6702 | 0.5926
| 28 | 2023-11-08 09:36 | 181.7578 | 0.97031 | 19.7797 | 35.3742 | 0.59224 | 42.041 | 1.8163e-05 | 10765 | 59.2272 | 100.0375 | 0.59227 | 0.0045688 | 19.7797 | 36.3445 | 42.042 | 355.5179 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.59339 | 0.19333 | 0.59224 | [6 empty] | 0.59379 | 0.26169 | 0.59224 | [6 empty] | 19.525 | 35.3856 | 50.6335 | 42.0942 | 1.8152e-05 | 14824 | 81.5591 | 272499.1596 | 5798.78 | 50.6335 | -0.56463 | 0.075798 | -0.042163 | -0.00069724 | 0.032938 | 5796.87 | 50.6502 | 272588.9445 | -0.56449 | 0.075794 | -0.042145 | -0.00069747 | 0.032952 | 5796.8692 | 50.6502 | 0.59224
| 29 | 2023-11-08 09:42 | 182.8343 | 0.97032 | 19.6947 | 35.3748 | 0.23507 | 42.055 | 1.8159e-05 | 4311 | 23.5787 | 100.2942 | 0.23579 | 0.30518 | 19.6947 | 36.3451 | 42.056 | 355.4636 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.23623 | 0.49449 | 0.23507 | 0.23505 | 0.4543 | 0.051273 | 0.15601 | 0.25 | 0.29468 | 0.23639 | 0.56308 | 0.23507 | 0.23505 | 0.52287 | 0.051295 | 0.15608 | 0.25 | 0.29472 | 19.4558 | 35.3903 | 20.0982 | 42.1105 | 1.8149e-05 | 5919 | 32.3736 | 108223.6224 | 5798.78 | 20.0982 | -0.96568 | 0.14959 | -0.16022 | -0.00027683 | -0.01091 | 5799.4127 | 20.096 | 108211.8162 | -0.96573 | 0.14961 | -0.16025 | -0.0002768 | -0.010918 | 5799.4131 | 20.096 | 0.23507
| 30 | 2023-11-08 09:45 | 182.9167 | 0.97037 | 19.6546 | 35.3634 | 0.23545 | 42.0481 | 1.8157e-05 | 4319 | 23.6118 | 100.2718 | 0.23612 | 0.28268 | 19.6546 | 36.3337 | 42.0491 | 355.4354 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.23656 | 0.47185 | 0.23545 | [6 empty] | 0.23673 | 0.54043 | 0.23545 | [6 empty] | 19.4033 | 35.3789 | 20.1332 | 42.1056 | 1.8146e-05 | 5932 | 32.4301 | 108414.8914 | 5798.78 | 20.1332 | -0.96491 | 0.14927 | -0.1598 | -0.00027737 | -0.010807 | 5799.4067 | 20.1311 | 108403.176 | -0.96496 | 0.14929 | -0.15983 | -0.00027734 | -0.010815 | 5799.4071 | 20.1311 | 0.23545
| 31 | 2023-11-08 09:49 | 181.7824 | 0.97041 | 19.5974 | 35.3596 | 0.23463 | 42.0528 | 1.8154e-05 | 4274 | 23.5116 | 100.1964 | 0.23512 | 0.20757 | 19.5974 | 36.33 | 42.0537 | 355.3979 | 1.4621 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.23556 | 0.39656 | 0.23463 | [6 empty] | 0.23572 | 0.4651 | 0.23463 | [6 empty] | 19.3503 | 35.3755 | 20.0608 | 42.11 | 1.8144e-05 | 5874 | 32.3134 | 108050.235 | 5798.78 | 20.0608 | -0.96637 | 0.14987 | -0.16069 | -0.00027642 | -0.011093 | 5799.4233 | 20.0586 | 108038.2504 | -0.96642 | 0.14989 | -0.16072 | -0.00027639 | -0.011101 | 5799.4237 | 20.0586 | 0.23463
| 32 | 2023-11-08 10:01 | 183.189 | 0.97059 | 20.5267 | 35.4897 | 5.4998 | 42.0583 | 1.8198e-05 | 100553 | 548.9032 | 99.8501 | 5.489 | -0.19646 | 20.5267 | 36.4603 | 42.0592 | 356.0222 | 1.4622 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 5.4995 | -0.0069932 | 5.4998 | 5.5256 | -0.036051 | 0.11973 | 0.36431 | 0.25 | 0.44184 | 5.5032 | 0.061038 | 5.4998 | 5.5256 | 0.032132 | 0.11975 | 0.36437 | 0.25 | 0.44189 | 20.2343 | 35.113 | 474.9494 | 41.668 | 1.8178e-05 | 140146 | 765.0352 | 2526601.9111 | 5798.78 | 474.9494 | 0.40254 | 0.056343 | -0.003182 | -0.0065229 | 0.046638 | 5796.0755 | 475.171 | 2527780.8277 | 0.40274 | 0.056321 | -0.0031804 | -0.0065259 | 0.046615 | 5796.0769 | 475.1709 | 5.4998
| 33 | 2023-11-08 10:09 | 182.4615 | 0.97071 | 19.6105 | 35.4841 | 5.5076 | 42.1955 | 1.8157e-05 | 100369 | 550.0831 | 99.9236 | 5.5008 | -0.12286 | 19.6105 | 36.4548 | 42.1965 | 355.4348 | 1.4623 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 5.5113 | 0.066479 | 5.5076 | [6 empty] | 5.515 | 0.13479 | 5.5076 | [6 empty] | 19.7463 | 35.1075 | 474.8286 | 41.7375 | 1.8156e-05 | 139554 | 764.8407 | 2533168.407 | 5798.78 | 474.8286 | 0.40366 | 0.056219 | -0.0031775 | -0.006533 | 0.046509 | 5796.0831 | 475.0495 | 2534347.1031 | 0.40387 | 0.056197 | -0.003176 | -0.0065361 | 0.046485 | 5796.0844 | 475.0494 | 5.5076
| 34 | 2023-11-08 10:15 | 183.9829 | 0.97077 | 19.3608 | 35.4801 | 5.5693 | 42.2302 | 1.8146e-05 | 102101 | 554.9484 | 99.6893 | 5.5495 | -0.35644 | 19.3608 | 36.4509 | 42.2312 | 355.2739 | 1.4624 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 5.56 | -0.16764 | 5.5693 | [6 empty] | 5.5638 | -0.099429 | 5.5693 | [6 empty] | 19.4852 | 35.0955 | 479.8476 | 41.7641 | 1.8144e-05 | 142205 | 772.9253 | 2563230.5233 | 5798.78 | 479.8476 | 0.40879 | 0.055654 | -0.0031406 | -0.0066085 | 0.045905 | 5796.1181 | 480.068 | 2564407.7148 | 0.40899 | 0.055632 | -0.0031391 | -0.0066116 | 0.045881 | 5796.1194 | 480.0679 | 5.5693
| 35 | 2023-11-08 10:23 | 64.2098 | 0.9709 | 19.2188 | 35.4801 | #NUM! | 42.2527 | 1.814e-05 | 5 | 0.07787 | NaN | 0.0007787 | #NUM! | 19.2188 | 36.451 | 42.2537 | 355.1829 | 1.4624 | 93.52 | 44.7 | 0.02235 | 1 | 0.99932 | 1.0026 | 0.00078017 | #NUM! | #NUM! | [6 empty] | 0.00078071 | #NUM! | #NUM! | [6 empty] | 19.344 | 35.4965 | 0 | 42.2521 | 1.8146e-05 | 0 | 0 | 0 | 5798.78 | 0 | #NUM! | #NUM! | #DIV/0! | 0 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM!
Target Type | Testing.TestPoints.TestPoint | Testing.TestPoints.TestPoint | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | [10 empty] | #VALUE! | #VALUE! | #VALUE! | [16 empty] | 20.0416 | 9.856 | 624.643 | 12.4721 | 1.7735e-05 | 182877 | 1006.1488 | 1019415.2668 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | [12 empty] | #DIV/0! | #DIV/0!
Target Attribute | Number | Timestamp | 10 | 7 | 24 | 21 | 34 | 48 | 94 | 26 | 18 | 19 | 16 | 31 | [10 empty] | 1.0023 | 16.0374 | -52.8311 | [16 empty] | 19.9756 | 10.0557 | 0 | 12.7054 | 1.7735e-05 | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0! | [12 empty] | #DIV/0! | #DIV/0!
Target Setup Z Order | -1 | -1 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89
Display Units | - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | %
Format |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
-0.0066582

## (TBL) Coriolis N2 60 bar
Table: Coriolis | [57 empty] | a0 | 0.08203
[58 empty] | a1 | -0.012922
Job Number | EMR013 | [49 empty] | 5778303.2906 | [6 empty] | a2 | -0.11908
Test Number | T03119 | [49 empty] | 719576.9504 | [6 empty] | a3 | -0.067147
Test Title | 3" Coriolis Intercomparison 60 Bar | [25 empty] | 0 | [30 empty] | a4 | 0.1215
Test Description | 3" Coriolis Intercomparison 60 Bar | [56 empty] | b0 | -0.00011761
[39 empty] | 0.18174 |  |  |  | 0.25021 | [14 empty] | b1 | -0.016601
[43 empty] | 0.40417 | [14 empty] | cp | -0.042815
ONLY PRESSURE CORRECTION APPLIED - NO COMPRESSIBILITY
[15 empty] | RefProp version 10 | [13 empty] | Averaged values | [8 empty] | Averaged values | [5 empty] | Reference Turbine Calcs
| Test Point | TestPoint Timestamp | Collection Time (t) | Ambient Pressure (Pamb) | Test Temperature (T) | Test Pressure (Pgauge) | Reference Mass Flow Rate (ṁ) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | K Factor (K) | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Test Temperature (T) | Test Pressure (Pabs) | Gas Density (ρ) | Speed of Sound | Isentropic expansion | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | Pressure correction fp | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Test Temperature (T) | Test Pressure (Pgauge) | Meter Est Volume Flow (Qi) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | Reynolds | K nominal | Meter Nominal Flow (Q) | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Meter Est Volume Flow (Qi) | Reynolds | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Reference Volume Flow Rate (Q) | Reference Mass Flow Rate (ṁ)
| - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | % | °C | bara | kg/m³ | m/s | - | Hz | mm | mm | - | - | - | kg/s | % | kg/s | kg/s | % | % | % | % | % | kg/s | % | kg/s | kg/s | % | % | % | % | % | °C | barg | m³/h | kg/m³ | Pa s | pulse | Hz | - | - | m3/s | - | % | % | % | % | pulse/m³ | m³/h | - | - | % | % | % | % | pulse/m³ | m³/h | kg/s
| 8 | 2023-11-07 12:36 | 184.3922 | 0.98073 | 20.0102 | 60.8959 | 12.9356 | 71.5122 | 1.8721e-05 | 238603 | 1293.9973 | 100.0377 | 12.94 | 0.03382 | 20.0102 | 61.8766 | 71.5137 | 362.3122 | 1.5172 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0046 | 12.9904 | 0.42381 | 12.9356 | [15 empty] | 20.2602 | 59.6654 | 664.9724 | 70.0275 | 1.8702e-05 | 197506 | 1071.1191 | 5778303.2906 | 5798.78 | 664.9724 | 0.7618 | 0.014312 | -0.0013326 | -0.009132 | 0.0038471 | 5798.5569 | 664.998 | 5778525.5989 | 0.76182 | 0.01431 | -0.0013326 | -0.0091324 | 0.0038452 | 5798.557 | 664.998 | 12.9356
| 11 | 2023-11-07 12:54 | 185.6111 | 0.98085 | 19.9818 | 60.7734 | 11.4351 | 71.3793 | 1.8717e-05 | 211842 | 1141.3219 | 99.8189 | 11.4132 | -0.19136 | 19.9818 | 61.7542 | 71.3808 | 362.2579 | 1.5169 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 11.4576 | 0.19678 | 11.4351 | 11.4318 | 0.28543 | 0.077903 | 0.15292 | 0.25 | 0.29306 | 11.465 | 0.26185 | 11.4351 | 11.4318 | 0.35056 | 0.077954 | 0.15302 | 0.25 | 0.29311 | 20.1323 | 59.8147 | 586.0681 | 70.2344 | 1.8701e-05 | 175221 | 944.0223 | 5108185.7711 | 5798.78 | 586.0681 | 0.70827 | 0.019858 | -0.0015151 | -0.0080524 | 0.010291 | 5798.1833 | 586.1284 | 5108711.5027 | 0.70831 | 0.019853 | -0.0015149 | -0.0080532 | 0.010285 | 5798.1836 | 586.1284 | 11.4351
| 12 | 2023-11-07 12:58 | 181.8373 | 0.98089 | 19.9866 | 60.7639 | 11.4339 | 71.3671 | 1.8717e-05 | 207855 | 1143.0822 | 99.9832 | 11.4308 | -0.027104 | 19.9866 | 61.7448 | 71.3686 | 362.2583 | 1.5169 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 11.4753 | 0.3616 | 11.4339 | [6 empty] | 11.4827 | 0.42678 | 11.4339 | [6 empty] | 20.1437 | 59.8056 | 586.1212 | 70.2208 | 1.8701e-05 | 171674 | 944.1077 | 5107588.3496 | 5798.78 | 586.1212 | 0.70822 | 0.019864 | -0.0015152 | -0.0080528 | 0.010296 | 5798.183 | 586.1815 | 5108114.2783 | 0.70826 | 0.019859 | -0.0015151 | -0.0080536 | 0.01029 | 5798.1833 | 586.1815 | 11.4339
| 13 | 2023-11-07 13:02 | 181.8128 | 0.98093 | 19.9759 | 60.7543 | 11.431 | 71.3591 | 1.8716e-05 | 207648 | 1142.0976 | 99.923 | 11.421 | -0.087343 | 19.9759 | 61.7353 | 71.3606 | 362.2485 | 1.5168 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 11.4654 | 0.30105 | 11.431 | [6 empty] | 11.4728 | 0.36619 | 11.431 | [6 empty] | 20.142 | 59.7961 | 586.0566 | 70.2103 | 1.87e-05 | 171632 | 944.0037 | 5106343.9416 | 5798.78 | 586.0566 | 0.70811 | 0.019876 | -0.0015156 | -0.0080519 | 0.010308 | 5798.1823 | 586.117 | 5106870.3586 | 0.70815 | 0.019871 | -0.0015155 | -0.0080528 | 0.010302 | 5798.1826 | 586.117 | 11.431
| 14 | 2023-11-07 13:06 | 301.8395 | 0.98094 | 19.9755 | 60.7465 | 11.4298 | 71.3502 | 1.8716e-05 | 344389 | 1140.9671 | 99.8342 | 11.4097 | -0.17612 | 19.9755 | 61.7275 | 71.3518 | 362.246 | 1.5168 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 11.454 | 0.21187 | 11.4298 | [6 empty] | 11.4615 | 0.27695 | 11.4298 | [6 empty] | 20.1421 | 59.7881 | 586.0738 | 70.2011 | 1.87e-05 | 284946 | 944.0314 | 5105876.2755 | 5798.78 | 586.0738 | 0.70807 | 0.01988 | -0.0015158 | -0.0080522 | 0.010312 | 5798.182 | 586.1342 | 5106402.8479 | 0.70812 | 0.019875 | -0.0015156 | -0.008053 | 0.010306 | 5798.1824 | 586.1342 | 11.4298
| 15 | 2023-11-07 13:12 | 301.7874 | 0.98095 | 19.9628 | 60.7368 | 11.4293 | 71.3425 | 1.8715e-05 | 344809 | 1142.5558 | 99.9777 | 11.4256 | -0.032623 | 19.9628 | 61.7177 | 71.3441 | 362.2349 | 1.5168 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 11.47 | 0.35584 | 11.4293 | [6 empty] | 11.4774 | 0.42103 | 11.4293 | [6 empty] | 20.1384 | 59.7782 | 586.1338 | 70.1908 | 1.87e-05 | 284926 | 944.1281 | 5105749.7722 | 5798.78 | 586.1338 | 0.70806 | 0.019881 | -0.0015158 | -0.0080531 | 0.010312 | 5798.182 | 586.1943 | 5106276.3407 | 0.7081 | 0.019876 | -0.0015157 | -0.0080539 | 0.010307 | 5798.1823 | 586.1942 | 11.4293
| 17 | 2023-11-07 13:33 | 181.7753 | 0.98093 | 19.8918 | 60.6243 | 9.8775 | 71.2327 | 1.871e-05 | 179183 | 985.7389 | 99.8145 | 9.8574 | -0.20396 | 19.8918 | 61.6052 | 71.2342 | 362.1557 | 1.5166 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 9.8956 | 0.18294 | 9.8775 | 9.8759 | 0.088906 | 0.080298 | 0.15762 | 0.25 | 0.29554 | 9.902 | 0.24804 | 9.8775 | 9.8759 | 0.15395 | 0.080349 | 0.15772 | 0.25 | 0.29559 | 20.0422 | 59.9122 | 505.2141 | 70.3713 | 1.8699e-05 | 147926 | 813.7848 | 4412375.1731 | 5798.78 | 505.2141 | 0.64467 | 0.027204 | -0.0017658 | -0.0069439 | 0.018494 | 5797.7076 | 505.3076 | 4413191.3561 | 0.64475 | 0.027194 | -0.0017655 | -0.0069451 | 0.018484 | 5797.7082 | 505.3075 | 9.8775
| 18 | 2023-11-07 13:37 | 181.8808 | 0.98093 | 19.8797 | 60.6153 | 9.8762 | 71.2257 | 1.8709e-05 | 178970 | 983.9963 | 99.6514 | 9.84 | -0.36701 | 19.8797 | 61.5962 | 71.2273 | 362.1452 | 1.5166 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 9.8781 | 0.019189 | 9.8762 | [6 empty] | 9.8845 | 0.084185 | 9.8762 | [6 empty] | 20.0318 | 59.9035 | 505.1977 | 70.3641 | 1.8698e-05 | 148007 | 813.7584 | 4411931.1153 | 5798.78 | 505.1977 | 0.64463 | 0.027209 | -0.0017661 | -0.0069439 | 0.018499 | 5797.7073 | 505.2911 | 4412747.4344 | 0.64471 | 0.027199 | -0.0017657 | -0.0069452 | 0.018489 | 5797.7079 | 505.2911 | 9.8762
| 19 | 2023-11-07 13:40 | 181.751 | 0.98097 | 19.8835 | 60.6107 | 9.8752 | 71.2194 | 1.8709e-05 | 179085 | 985.3314 | 99.7972 | 9.8533 | -0.22126 | 19.8835 | 61.5917 | 71.221 | 362.1463 | 1.5165 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 9.8915 | 0.16547 | 9.8752 | [6 empty] | 9.8979 | 0.23056 | 9.8752 | [6 empty] | 20.0283 | 59.8992 | 505.1723 | 70.3602 | 1.8698e-05 | 147894 | 813.7175 | 4411520.4107 | 5798.78 | 505.1723 | 0.64459 | 0.027214 | -0.0017663 | -0.0069437 | 0.018504 | 5797.707 | 505.2658 | 4412336.8706 | 0.64467 | 0.027204 | -0.0017659 | -0.006945 | 0.018493 | 5797.7076 | 505.2657 | 9.8752
| 20 | 2023-11-07 13:44 | 301.746 | 0.981 | 19.874 | 60.6036 | 9.8758 | 71.2139 | 1.8708e-05 | 297032 | 984.3775 | 99.6945 | 9.8438 | -0.32396 | 19.874 | 61.5846 | 71.2155 | 362.1381 | 1.5165 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 9.8819 | 0.062311 | 9.8758 | [6 empty] | 9.8883 | 0.12734 | 9.8758 | [6 empty] | 20.0188 | 59.8917 | 505.2463 | 70.3542 | 1.8698e-05 | 245572 | 813.8367 | 4411924.2894 | 5798.78 | 505.2463 | 0.64463 | 0.027209 | -0.0017661 | -0.0069449 | 0.018498 | 5797.7073 | 505.3398 | 4412740.5624 | 0.64471 | 0.0272 | -0.0017658 | -0.0069462 | 0.018488 | 5797.7079 | 505.3398 | 9.8758
| 21 | 2023-11-07 13:50 | 301.777 | 0.98101 | 19.8762 | 60.5986 | 9.8747 | 71.2075 | 1.8708e-05 | 296888 | 983.7993 | 99.647 | 9.838 | -0.37143 | 19.8762 | 61.5796 | 71.2091 | 362.1381 | 1.5165 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 9.8761 | 0.01462 | 9.8747 | [6 empty] | 9.8825 | 0.079615 | 9.8747 | [6 empty] | 20.014 | 59.887 | 505.2191 | 70.3501 | 1.8697e-05 | 245584 | 813.793 | 4411504.6801 | 5798.78 | 505.2191 | 0.64459 | 0.027214 | -0.0017663 | -0.0069447 | 0.018503 | 5797.707 | 505.3126 | 4412321.0961 | 0.64467 | 0.027204 | -0.001766 | -0.006946 | 0.018493 | 5797.7077 | 505.3126 | 9.8747
| 23 | 2023-11-07 14:31 | 182.7249 | 0.98121 | 19.925 | 60.4931 | 8.0219 | 71.0728 | 1.8708e-05 | 146127 | 799.7106 | 99.7216 | 7.9971 | -0.30877 | 19.925 | 61.4744 | 71.0744 | 362.1393 | 1.5162 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 8.028 | 0.076705 | 8.0219 | 8.022 | 0.15809 | 0.1043 | 0.31736 | 0.25 | 0.404 | 8.0332 | 0.14174 | 8.0219 | 8.022 | 0.22318 | 0.10437 | 0.31757 | 0.25 | 0.40417 | 20.0044 | 60.0265 | 409.424 | 70.5136 | 1.87e-05 | 120505 | 659.4888 | 3582790.7198 | 5798.78 | 409.424 | 0.55422 | 0.038323 | -0.002199 | -0.0056281 | 0.030496 | 5797.0116 | 409.5489 | 3583883.6631 | 0.55435 | 0.038307 | -0.0021983 | -0.0056298 | 0.030479 | 5797.0126 | 409.5488 | 8.0219
| 24 | 2023-11-07 14:37 | 181.9228 | 0.98123 | 19.9372 | 60.4909 | 8.0222 | 71.0669 | 1.8708e-05 | 145557 | 800.1031 | 99.7667 | 8.001 | -0.26373 | 19.9372 | 61.4721 | 71.0684 | 362.1467 | 1.5162 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 8.032 | 0.1219 | 8.0222 | [6 empty] | 8.0372 | 0.18696 | 8.0222 | [6 empty] | 20.011 | 60.0239 | 409.4682 | 70.5087 | 1.87e-05 | 119989 | 659.56 | 3582890.52 | 5798.78 | 409.4682 | 0.55423 | 0.038322 | -0.0021989 | -0.0056285 | 0.030494 | 5797.0117 | 409.5931 | 3583983.427 | 0.55437 | 0.038305 | -0.0021982 | -0.0056303 | 0.030477 | 5797.0127 | 409.593 | 8.0222
| 25 | 2023-11-07 14:41 | 181.7673 | 0.98124 | 19.9219 | 60.485 | 8.0218 | 71.0643 | 1.8708e-05 | 145649 | 801.2937 | 99.9199 | 8.0129 | -0.1105 | 19.9219 | 61.4662 | 71.0659 | 362.135 | 1.5162 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 8.0439 | 0.27567 | 8.0218 | [6 empty] | 8.0491 | 0.34084 | 8.0218 | [6 empty] | 19.9974 | 60.0184 | 409.4633 | 70.5062 | 1.87e-05 | 119885 | 659.552 | 3582848.8955 | 5798.78 | 409.4633 | 0.55423 | 0.038322 | -0.002199 | -0.0056288 | 0.030494 | 5797.0117 | 409.5882 | 3583941.8 | 0.55436 | 0.038306 | -0.0021983 | -0.0056305 | 0.030477 | 5797.0127 | 409.5881 | 8.0218
| 27 | 2023-11-07 14:53 | 181.755 | 0.98121 | 19.9041 | 60.4332 | 6.537 | 71.0095 | 1.8706e-05 | 118670 | 652.912 | 99.9216 | 6.5291 | -0.12021 | 19.9041 | 61.4144 | 71.0111 | 362.1083 | 1.5161 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5543 | 0.26552 | 6.537 | 6.5356 | 0.16598 | 0.092878 | 0.18231 | 0.25 | 0.30942 | 6.5586 | 0.33069 | 6.537 | 6.5356 | 0.23109 | 0.092936 | 0.18243 | 0.25 | 0.30948 | 19.9601 | 60.1262 | 332.9999 | 70.6405 | 1.8701e-05 | 97491 | 536.387 | 2919182.4271 | 5798.78 | 332.9999 | 0.46526 | 0.049171 | -0.0027352 | -0.0045783 | 0.041858 | 5796.3528 | 333.1394 | 2920404.8367 | 0.46544 | 0.049149 | -0.0027339 | -0.0045802 | 0.041835 | 5796.3541 | 333.1393 | 6.537
| 28 | 2023-11-07 15:03 | 181.7596 | 0.98129 | 19.8853 | 60.4131 | 6.5351 | 70.9917 | 1.8704e-05 | 118487 | 651.8885 | 99.7942 | 6.5189 | -0.24759 | 19.8853 | 61.3944 | 70.9932 | 362.0902 | 1.5161 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.544 | 0.13748 | 6.5351 | [6 empty] | 6.5483 | 0.20258 | 6.5351 | [6 empty] | 19.9361 | 60.1065 | 332.9777 | 70.6245 | 1.8699e-05 | 97487 | 536.3513 | 2918552.4155 | 5798.78 | 332.9777 | 0.46517 | 0.049182 | -0.002736 | -0.0045784 | 0.041868 | 5796.3522 | 333.1172 | 2919774.8579 | 0.46535 | 0.049161 | -0.0027347 | -0.0045803 | 0.041845 | 5796.3535 | 333.1171 | 6.5351
| 29 | 2023-11-07 15:06 | 181.7868 | 0.98131 | 19.8825 | 60.4079 | 6.5363 | 70.9865 | 1.8704e-05 | 118464 | 651.6645 | 99.7416 | 6.5166 | -0.30012 | 19.8825 | 61.3892 | 70.988 | 362.0868 | 1.5161 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5418 | 0.084704 | 6.5363 | [6 empty] | 6.5461 | 0.14976 | 6.5363 | [6 empty] | 19.9294 | 60.1014 | 333.0577 | 70.6205 | 1.8699e-05 | 97525 | 536.4801 | 2919148.2131 | 5798.78 | 333.0577 | 0.46526 | 0.049172 | -0.0027354 | -0.0045796 | 0.041857 | 5796.3528 | 333.1972 | 2920370.5788 | 0.46544 | 0.04915 | -0.0027342 | -0.0045815 | 0.041834 | 5796.3541 | 333.1971 | 6.5363
| 30 | 2023-11-07 15:10 | 181.7812 | 0.98132 | 19.8774 | 60.4029 | 6.535 | 70.9822 | 1.8704e-05 | 118650 | 652.7078 | 99.92 | 6.5271 | -0.12184 | 19.8774 | 61.3843 | 70.9838 | 362.0821 | 1.5161 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5523 | 0.26363 | 6.535 | [6 empty] | 6.5565 | 0.32881 | 6.535 | [6 empty] | 19.925 | 60.0964 | 333.0168 | 70.616 | 1.8699e-05 | 97510 | 536.4142 | 2918649.4288 | 5798.78 | 333.0168 | 0.46518 | 0.04918 | -0.002736 | -0.0045791 | 0.041865 | 5796.3523 | 333.1562 | 2919871.8407 | 0.46536 | 0.049159 | -0.0027347 | -0.0045811 | 0.041843 | 5796.3536 | 333.1562 | 6.535
| 31 | 2023-11-07 15:14 | 181.7712 | 0.9813 | 19.8764 | 60.3988 | 6.5348 | 70.9777 | 1.8704e-05 | 118421 | 651.4838 | 99.7356 | 6.5148 | -0.30618 | 19.8764 | 61.3801 | 70.9793 | 362.0802 | 1.516 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.54 | 0.078553 | 6.5348 | [6 empty] | 6.5442 | 0.14361 | 6.5348 | [6 empty] | 19.9195 | 60.0924 | 333.0214 | 70.6129 | 1.8698e-05 | 97506 | 536.4216 | 2918612.4357 | 5798.78 | 333.0214 | 0.46518 | 0.049181 | -0.002736 | -0.0045793 | 0.041866 | 5796.3523 | 333.1609 | 2919834.8441 | 0.46536 | 0.049159 | -0.0027348 | -0.0045812 | 0.041843 | 5796.3536 | 333.1608 | 6.5348
| 33 | 2023-11-07 15:25 | 181.8083 | 0.98138 | 19.859 | 60.3532 | 4.8033 | 70.93 | 1.8702e-05 | 87037 | 478.7295 | 99.7236 | 4.7873 | -0.33272 | 19.859 | 61.3346 | 70.9316 | 362.0556 | 1.516 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.8058 | 0.051551 | 4.8033 | 4.8027 | 0.08621 | 0.041914 | 0.12753 | 0.25 | 0.28065 | 4.8089 | 0.1166 | 4.8033 | 4.8027 | 0.15128 | 0.041942 | 0.12762 | 0.25 | 0.28069 | 19.8879 | 60.1923 | 244.3147 | 70.7367 | 1.8699e-05 | 71548 | 393.5354 | 2144819.3947 | 5798.78 | 244.3147 | 0.33139 | 0.063692 | -0.0038191 | -0.0033599 | 0.056513 | 5795.5029 | 244.4529 | 2146032.1817 | 0.33164 | 0.063668 | -0.0038167 | -0.0033618 | 0.05649 | 5795.5043 | 244.4528 | 4.8033
| 34 | 2023-11-07 15:30 | 181.7998 | 0.98143 | 19.8524 | 60.3435 | 4.8022 | 70.9208 | 1.8701e-05 | 87085 | 479.0158 | 99.8047 | 4.7902 | -0.25171 | 19.8524 | 61.3249 | 70.9224 | 362.0485 | 1.5159 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.8086 | 0.13279 | 4.8022 | [6 empty] | 4.8117 | 0.1979 | 4.8022 | [6 empty] | 19.8796 | 60.1828 | 244.292 | 70.7281 | 1.8699e-05 | 71538 | 393.4987 | 2144423.0123 | 5798.78 | 244.292 | 0.33131 | 0.0637 | -0.0038199 | -0.0033597 | 0.05652 | 5795.5025 | 244.4301 | 2145635.727 | 0.33156 | 0.063676 | -0.0038176 | -0.0033616 | 0.056497 | 5795.5039 | 244.4301 | 4.8022
| 35 | 2023-11-07 15:34 | 181.8033 | 0.98146 | 19.853 | 60.3373 | 4.8025 | 70.9135 | 1.8701e-05 | 87041 | 478.7648 | 99.7464 | 4.7876 | -0.30994 | 19.853 | 61.3188 | 70.9151 | 362.0471 | 1.5159 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.8061 | 0.074285 | 4.8025 | [6 empty] | 4.8092 | 0.13935 | 4.8025 | [6 empty] | 19.8758 | 60.1772 | 244.3249 | 70.7228 | 1.8699e-05 | 71549 | 393.5518 | 2144583.8687 | 5798.78 | 244.3249 | 0.33134 | 0.063697 | -0.0038196 | -0.0033602 | 0.056517 | 5795.5027 | 244.4631 | 2145796.6022 | 0.33159 | 0.063673 | -0.0038173 | -0.0033621 | 0.056494 | 5795.5041 | 244.463 | 4.8025
| 37 | 2023-11-07 15:46 | 181.7594 | 0.98153 | 19.8206 | 60.2996 | 3.2867 | 70.8791 | 1.8699e-05 | 59578 | 327.785 | 99.7987 | 3.2778 | -0.26943 | 19.8206 | 61.2811 | 70.8806 | 362.015 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2905 | 0.11465 | 3.2867 | 3.2867 | 0.11077 | 0.0033856 | 0.010301 | 0.25 | 0.25021 | 3.2926 | 0.17975 | 3.2867 | 3.2867 | 0.17587 | 0.0033855 | 0.010301 | 0.25 | 0.25021 | 19.8601 | 60.2316 | 167.0304 | 70.7899 | 1.8699e-05 | 48902 | 269.048 | 1467465.1165 | 5798.78 | 167.0304 | 0.16657 | 0.076357 | -0.0058261 | -0.0022973 | 0.068234 | 5794.8233 | 167.1445 | 1468467.1085 | 0.16686 | 0.076341 | -0.0058216 | -0.0022989 | 0.06822 | 5794.8241 | 167.1445 | 3.2867
| 38 | 2023-11-07 15:50 | 181.766 | 0.98155 | 19.8004 | 60.2881 | 3.2867 | 70.8715 | 1.8698e-05 | 59577 | 327.7675 | 99.7925 | 3.2777 | -0.27555 | 19.8004 | 61.2697 | 70.8731 | 361.9985 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2903 | 0.10841 | 3.2867 | [6 empty] | 3.2924 | 0.17352 | 3.2867 | [6 empty] | 19.8426 | 60.2201 | 167.0517 | 70.7815 | 1.8698e-05 | 48910 | 269.0822 | 1467554.4916 | 5798.78 | 167.0517 | 0.16659 | 0.076356 | -0.0058259 | -0.0022978 | 0.068232 | 5794.8234 | 167.1657 | 1468556.5185 | 0.16689 | 0.076339 | -0.0058214 | -0.0022993 | 0.068218 | 5794.8242 | 167.1657 | 3.2867
| 39 | 2023-11-07 15:54 | 181.7721 | 0.98155 | 19.8224 | 60.2874 | 3.2866 | 70.8646 | 1.8699e-05 | 59578 | 327.7621 | 99.7934 | 3.2776 | -0.27471 | 19.8224 | 61.269 | 70.8662 | 362.0126 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2902 | 0.10925 | 3.2866 | [6 empty] | 3.2924 | 0.17435 | 3.2866 | [6 empty] | 19.8326 | 60.2195 | 167.0427 | 70.7835 | 1.8698e-05 | 48909 | 269.0677 | 1467550.9496 | 5798.78 | 167.0427 | 0.16659 | 0.076356 | -0.0058261 | -0.0022977 | 0.068232 | 5794.8234 | 167.1567 | 1468552.9727 | 0.16689 | 0.076339 | -0.0058216 | -0.0022993 | 0.068218 | 5794.8242 | 167.1567 | 3.2866
| 40 | 2023-11-07 16:08 | 31.7656 | 0.98145 | 19.7643 | 60.2588 | 1.6131 | 70.8475 | 1.8696e-05 | 5111 | 160.8975 | 99.8091 | 1.609 | -0.25781 | 19.7643 | 61.2402 | 70.8491 | 361.9663 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6152 | 0.12598 | 1.6131 | [6 empty] | 1.6162 | 0.19111 | 1.6131 | [6 empty] | 19.7863 | 60.2536 | 81.9277 | 70.8355 | 1.8697e-05 | 4192 | 131.9668 | 720345.1935 | 5798.78 | 81.9277 | -0.14246 | 0.081699 | -0.013504 | -0.0011271 | 0.067068 | 5794.8909 | 81.9827 | 720828.6385 | -0.14217 | 0.081703 | -0.013492 | -0.0011279 | 0.067083 | 5794.89 | 81.9827 | 1.6131
| 41 | 2023-11-07 16:10 | 181.8069 | 0.98146 | 19.8072 | 60.2588 | 1.6131 | 70.8357 | 1.8698e-05 | 29240 | 160.83 | 99.7712 | 1.6083 | -0.29576 | 19.8072 | 61.2402 | 70.8373 | 361.9944 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6145 | 0.087894 | 1.6131 | [6 empty] | 1.6155 | 0.15299 | 1.6131 | [6 empty] | 19.7928 | 60.254 | 81.926 | 70.8341 | 1.8697e-05 | 23992 | 131.9642 | 720306.2517 | 5798.78 | 81.926 | -0.14248 | 0.081698 | -0.013504 | -0.0011271 | 0.067067 | 5794.8909 | 81.981 | 720789.6639 | -0.14219 | 0.081703 | -0.013493 | -0.0011279 | 0.067082 | 5794.8901 | 81.981 | 1.6131
| 43 | 2023-11-07 16:35 | 181.769 | 0.98157 | 19.8395 | 60.325 | 6.5289 | 70.9032 | 1.87e-05 | 118211 | 650.3365 | 99.6505 | 6.5034 | -0.39123 | 19.8395 | 61.3066 | 70.9048 | 362.0347 | 1.5159 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5284 | -0.0074118 | 6.5289 | 6.5293 | 0.0057405 | 0.015688 | 0.047733 | 0.25 | 0.25452 | 6.5327 | 0.057607 | 6.5289 | 6.5293 | 0.070764 | 0.015696 | 0.047759 | 0.25 | 0.25452 | 19.8317 | 60.0191 | 333.0016 | 70.5529 | 1.8693e-05 | 97499 | 536.3897 | 2916785.1269 | 5798.78 | 333.0016 | 0.4649 | 0.049213 | -0.0027385 | -0.0045806 | 0.041894 | 5796.3507 | 333.1411 | 2918007.5985 | 0.46509 | 0.049192 | -0.0027373 | -0.0045825 | 0.041872 | 5796.3519 | 333.1411 | 6.5289
| 44 | 2023-11-07 16:39 | 181.7881 | 0.98154 | 19.8554 | 60.3336 | 6.53 | 70.9087 | 1.8701e-05 | 118280 | 650.6477 | 99.6808 | 6.5065 | -0.3609 | 19.8554 | 61.3151 | 70.9102 | 362.0476 | 1.5159 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5316 | 0.023104 | 6.53 | [6 empty] | 6.5358 | 0.088138 | 6.53 | [6 empty] | 19.8479 | 60.0275 | 333.0348 | 70.5581 | 1.8694e-05 | 97519 | 536.4433 | 2917158.8778 | 5798.78 | 333.0348 | 0.46496 | 0.049207 | -0.002738 | -0.0045808 | 0.041888 | 5796.351 | 333.1744 | 2918381.3233 | 0.46514 | 0.049185 | -0.0027368 | -0.0045827 | 0.041866 | 5796.3523 | 333.1743 | 6.53
| 45 | 2023-11-07 16:43 | 181.8437 | 0.98158 | 19.8608 | 60.3383 | 6.5291 | 70.9127 | 1.8702e-05 | 118273 | 650.4101 | 99.6593 | 6.5041 | -0.38243 | 19.8608 | 61.3199 | 70.9142 | 362.0526 | 1.5159 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 6.5292 | 0.0015292 | 6.5291 | [6 empty] | 6.5334 | 0.066547 | 6.5291 | [6 empty] | 19.8561 | 60.0321 | 332.9705 | 70.5612 | 1.8694e-05 | 97530 | 536.3397 | 2916654.4067 | 5798.78 | 332.9705 | 0.46488 | 0.049216 | -0.0027385 | -0.0045798 | 0.041897 | 5796.3505 | 333.1101 | 2917876.9174 | 0.46507 | 0.049194 | -0.0027373 | -0.0045817 | 0.041875 | 5796.3518 | 333.11 | 6.5291
| 46 | 2023-11-07 16:51 | 181.7552 | 0.98165 | 19.8274 | 60.3022 | 4.7905 | 70.8803 | 1.8699e-05 | 86792 | 477.5214 | 99.7381 | 4.7752 | -0.31838 | 19.8274 | 61.2838 | 70.8819 | 362.0202 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.7936 | 0.065531 | 4.7905 | 4.7899 | 0.094323 | 0.038526 | 0.11722 | 0.25 | 0.27612 | 4.7967 | 0.1306 | 4.7905 | 4.7899 | 0.15942 | 0.038555 | 0.11731 | 0.25 | 0.27616 | 19.8446 | 60.1432 | 243.8157 | 70.6924 | 1.8696e-05 | 71381 | 392.7315 | 2139429.482 | 5798.78 | 243.8157 | 0.3303 | 0.063797 | -0.00383 | -0.0033536 | 0.056613 | 5795.4971 | 243.9538 | 2140641.3721 | 0.33054 | 0.063773 | -0.0038277 | -0.0033555 | 0.05659 | 5795.4985 | 243.9537 | 4.7905
| 47 | 2023-11-07 16:55 | 181.778 | 0.9817 | 19.7814 | 60.2869 | 4.7901 | 70.8755 | 1.8697e-05 | 86859 | 477.8302 | 99.8105 | 4.7783 | -0.24597 | 19.7814 | 61.2686 | 70.877 | 361.9858 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.7967 | 0.13809 | 4.7901 | [6 empty] | 4.7998 | 0.20322 | 4.7901 | [6 empty] | 19.8203 | 60.1281 | 243.833 | 70.6818 | 1.8695e-05 | 71395 | 392.7594 | 2139413.6646 | 5798.78 | 243.833 | 0.33029 | 0.063797 | -0.0038303 | -0.0033542 | 0.056613 | 5795.4971 | 243.9711 | 2140625.5343 | 0.33054 | 0.063774 | -0.0038279 | -0.0033561 | 0.05659 | 5795.4985 | 243.9711 | 4.7901
| 48 | 2023-11-07 16:59 | 181.7647 | 0.98172 | 19.7717 | 60.2806 | 4.7893 | 70.871 | 1.8697e-05 | 86787 | 477.469 | 99.752 | 4.7747 | -0.30444 | 19.7717 | 61.2624 | 70.8725 | 361.9776 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 4.7931 | 0.07935 | 4.7893 | [6 empty] | 4.7962 | 0.14445 | 4.7893 | [6 empty] | 19.8065 | 60.1219 | 243.803 | 70.6784 | 1.8694e-05 | 71381 | 392.7111 | 2139131.0121 | 5798.78 | 243.803 | 0.33024 | 0.063803 | -0.003831 | -0.003354 | 0.056618 | 5795.4969 | 243.9411 | 2140342.8298 | 0.33048 | 0.063779 | -0.0038286 | -0.0033559 | 0.056595 | 5795.4982 | 243.941 | 4.7893
| 49 | 2023-11-07 17:06 | 181.7982 | 0.98176 | 19.8128 | 60.2656 | 3.2864 | 70.8424 | 1.8698e-05 | 59572 | 327.6821 | 99.7763 | 3.2768 | -0.29177 | 19.8128 | 61.2474 | 70.8439 | 362.0001 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2894 | 0.091957 | 3.2864 | 3.286 | 0.095244 | 0.0041478 | 0.01262 | 0.25 | 0.25032 | 3.2916 | 0.15705 | 3.2864 | 3.286 | 0.16035 | 0.0041576 | 0.01265 | 0.25 | 0.25032 | 19.7725 | 60.1979 | 167.0495 | 70.7755 | 1.8695e-05 | 48918 | 269.0786 | 1467675.6385 | 5798.78 | 167.0495 | 0.16663 | 0.076354 | -0.0058265 | -0.0022984 | 0.068229 | 5794.8236 | 167.1635 | 1468677.7012 | 0.16693 | 0.076337 | -0.005822 | -0.0022999 | 0.068215 | 5794.8243 | 167.1635 | 3.2864
| 50 | 2023-11-07 17:10 | 181.8131 | 0.98181 | 19.8023 | 60.2576 | 3.2856 | 70.8361 | 1.8697e-05 | 59563 | 327.6056 | 99.7783 | 3.2761 | -0.2898 | 19.8023 | 61.2394 | 70.8376 | 361.991 | 1.5158 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2887 | 0.09387 | 3.2856 | [6 empty] | 3.2908 | 0.15897 | 3.2856 | [6 empty] | 19.7574 | 60.19 | 167.0186 | 70.7706 | 1.8694e-05 | 48913 | 269.029 | 1467366.9614 | 5798.78 | 167.0186 | 0.16654 | 0.076359 | -0.0058281 | -0.0022981 | 0.068233 | 5794.8233 | 167.1327 | 1468368.8689 | 0.16684 | 0.076342 | -0.0058236 | -0.0022996 | 0.068219 | 5794.8241 | 167.1326 | 3.2856
| 51 | 2023-11-07 17:14 | 181.7899 | 0.98182 | 19.7639 | 60.2454 | 3.2861 | 70.8326 | 1.8696e-05 | 59568 | 327.675 | 99.7844 | 3.2767 | -0.28368 | 19.7639 | 61.2272 | 70.8342 | 361.9623 | 1.5157 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 3.2894 | 0.099904 | 3.2861 | [6 empty] | 3.2915 | 0.16502 | 3.2861 | [6 empty] | 19.7432 | 60.178 | 167.0673 | 70.7606 | 1.8693e-05 | 48921 | 269.1074 | 1467655.2974 | 5798.78 | 167.0673 | 0.16662 | 0.076354 | -0.005827 | -0.0022989 | 0.068228 | 5794.8236 | 167.1814 | 1468657.3361 | 0.16692 | 0.076338 | -0.0058225 | -0.0023004 | 0.068215 | 5794.8244 | 167.1813 | 3.2861
| 52 | 2023-11-07 17:22 | 181.7896 | 0.98182 | 19.7698 | 60.2258 | 1.6125 | 70.8084 | 1.8695e-05 | 29222 | 160.7463 | 99.7536 | 1.6075 | -0.31331 | 19.7698 | 61.2076 | 70.81 | 361.9604 | 1.5157 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6136 | 0.070007 | 1.6125 | 1.6118 | 0.075806 | 0.016203 | 0.031805 | 0.25 | 0.25201 | 1.6147 | 0.1351 | 1.6125 | 1.6118 | 0.14092 | 0.016213 | 0.031825 | 0.25 | 0.25202 | 19.7304 | 60.2219 | 81.9202 | 70.8148 | 1.8694e-05 | 23988 | 131.9548 | 720185.0346 | 5798.78 | 81.9202 | -0.14256 | 0.081697 | -0.013509 | -0.0011273 | 0.067061 | 5794.8913 | 81.9752 | 720668.3207 | -0.14226 | 0.081702 | -0.013498 | -0.001128 | 0.067076 | 5794.8904 | 81.9752 | 1.6125
| 53 | 2023-11-07 17:26 | 181.7512 | 0.9818 | 19.7606 | 60.2234 | 1.6129 | 70.8082 | 1.8695e-05 | 29229 | 160.8188 | 99.7755 | 1.6082 | -0.2914 | 19.7606 | 61.2052 | 70.8098 | 361.9537 | 1.5157 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6144 | 0.091987 | 1.6129 | [6 empty] | 1.6154 | 0.1571 | 1.6129 | [6 empty] | 19.7015 | 60.2187 | 81.9341 | 70.8191 | 1.8692e-05 | 23987 | 131.9771 | 720399.7915 | 5798.78 | 81.9341 | -0.14243 | 0.081699 | -0.013505 | -0.0011276 | 0.067066 | 5794.891 | 81.9891 | 720883.2604 | -0.14214 | 0.081704 | -0.013494 | -0.0011284 | 0.067081 | 5794.8901 | 81.9891 | 1.6129
| 54 | 2023-11-07 17:30 | 181.7755 | 0.98182 | 19.7018 | 60.2041 | 1.6118 | 70.8022 | 1.8692e-05 | 29204 | 160.6598 | 99.7437 | 1.6066 | -0.3232 | 19.7018 | 61.186 | 70.8038 | 361.9097 | 1.5157 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6128 | 0.059893 | 1.6118 | [6 empty] | 1.6138 | 0.125 | 1.6118 | [6 empty] | 19.6953 | 60.1996 | 81.9026 | 70.7989 | 1.8692e-05 | 23981 | 131.9265 | 719944.1169 | 5798.78 | 81.9026 | -0.1427 | 0.081695 | -0.013516 | -0.0011272 | 0.067052 | 5794.8918 | 81.9576 | 720427.1779 | -0.14241 | 0.0817 | -0.013505 | -0.001128 | 0.067067 | 5794.8909 | 81.9576 | 1.6118
| 55 | 2023-11-07 17:47 | 181.7976 | 0.98173 | 19.7341 | 60.1896 | 1.6108 | 70.7765 | 1.8693e-05 | 29191 | 160.5686 | 99.7469 | 1.6057 | -0.32002 | 19.7341 | 61.1713 | 70.778 | 361.9266 | 1.5156 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6119 | 0.062982 | 1.6108 | [6 empty] | 1.6129 | 0.12809 | 1.6108 | [6 empty] | 19.6627 | 60.1857 | 81.8619 | 70.7917 | 1.869e-05 | 23972 | 131.8609 | 719576.9504 | 5798.78 | 81.8619 | -0.14292 | 0.081692 | -0.013525 | -0.0011268 | 0.067039 | 5794.8925 | 81.9168 | 720059.6733 | -0.14263 | 0.081696 | -0.013514 | -0.0011275 | 0.067054 | 5794.8917 | 81.9168 | 1.6108
| 56 | 2023-11-07 17:51 | 181.7928 | 0.98171 | 19.699 | 60.1829 | 1.6111 | 70.7784 | 1.8691e-05 | 29204 | 160.6444 | 99.778 | 1.6064 | -0.2889 | 19.699 | 61.1646 | 70.78 | 361.9017 | 1.5156 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 1.6126 | 0.094162 | 1.6111 | [6 empty] | 1.6137 | 0.15929 | 1.6111 | [6 empty] | 19.6417 | 60.1786 | 81.8777 | 70.7893 | 1.8689e-05 | 23976 | 131.8864 | 719731.159 | 5798.78 | 81.8777 | -0.14283 | 0.081693 | -0.013523 | -0.0011271 | 0.067043 | 5794.8923 | 81.9327 | 720214.0137 | -0.14254 | 0.081698 | -0.013511 | -0.0011279 | 0.067058 | 5794.8914 | 81.9327 | 1.6111
| 57 | 2023-11-07 18:15 | 61.7871 | 0.98174 | 19.5826 | 60.1746 | #NUM! | 70.801 | 1.8686e-05 | 0 | 0 | NaN | 0 | #NUM! | 19.5826 | 61.1563 | 70.8025 | 361.8231 | 1.5157 | 92.97 | 44.7 | 0.02235 | 1 | 0.99935 | 1.0045 | 0 |  | #NUM! | [6 empty] | 0 | #NUM! | #NUM! | [6 empty] | 19.5555 | 60.1927 | 0 | 70.8294 | 1.8686e-05 | 0 | 0 | 0 | 5798.78 | 0 | #NUM! | #NUM! | #DIV/0! | 0 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM!
(End of Results Table)
Target Type | Testing.TestPoints.TestPoint | Testing.TestPoints.TestPoint | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc
Target Attribute | Number | Timestamp | 10 | 7 | 24 | 21 | 34 | 48 | 94 | 26 | 18 | 19 | 16 | 31
Target Setup Z Order | -1 | -1 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89
Display Units | - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | %
Format |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## (TBL) Coriolis CO2 27 bar
Table: Coriolis | [57 empty] | a0 | 0.08203
[58 empty] | a1 | -0.012922
Job Number | EMR013 | [49 empty] | 7398379.8676 | [6 empty] | a2 | -0.11908
Test Number | T03133 | [49 empty] | 872255.9613 | [6 empty] | a3 | -0.067147
Test Title | 3" Coriolis Intercomparison 27 Bar CO2 | [25 empty] | 0 | [30 empty] | a4 | 0.1215
[58 empty] | b0 | -0.00011761
[39 empty] | 0.28579 | [18 empty] | b1 | -0.016601
[58 empty] | cp | -0.042815
Test Description | 3" Coriolis Intercomparison 27 Bar CO2 | [33 empty] | ONLY PRESSURE CORRECTION APPLIED - NO COMPRESSIBILITY
[15 empty] | RefProp version 10 | [13 empty] | Averaged values | [8 empty] | Averaged values | [5 empty] | Reference Turbine Calcs
| Test Point | TestPoint Timestamp | Collection Time (t) | Ambient Pressure (Pamb) | Test Temperature (T) | Test Pressure (Pgauge) | Reference Mass Flow Rate (ṁ) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | K Factor (K) | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Test Temperature (T) | Test Pressure (Pabs) | Gas Density (ρ) | Speed of Sound | Isentropic expansion | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | Pressure correction fp | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Test Temperature (T) | Test Pressure (Pgauge) | Reference Volume Flow Rate (Q) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | Reynolds | K nominal | Meter Nominal Flow (Q) | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Meter Est Volume Flow (Qi) | Reynolds | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Reference Volume Flow Rate (Q) | Reference Mass Flow Rate (ṁ)
| - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | % | °C | bara | kg/m³ | m/s | - | Hz | mm | mm | - | - | - | kg/s | % | kg/s | kg/s | % | % | % | % | % | kg/s | % | kg/s | kg/s | % | % | % | % | % | °C | barg | m³/h | kg/m³ | Pa s | pulse | Hz | - | - | m3/s | - | % | % | % | % | pulse/m³ | m³/h | - | - | % | % | % | % | pulse/m³ | m³/h | kg/s
| 2 | 2023-11-09 08:40 | 182.296 | 0.96676 | 19.9119 | 27.8303 | 13.4046 | 62.8812 | 1.5178e-05 | 277691 | 1523.2969 | 113.621 | 15.233 | 13.6396 | 19.9119 | 28.797 | 62.8835 | 241.1426 | 1.2698 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 15.2408 | 13.6982 | 13.4046 |  |  | 3.6887 |  |  |  | 15.2632 | 13.8654 | 13.4046 |  |  | 3.6941 |  |  |  | 20.1689 | 26.2071 | 825.7838 | 58.447 | 1.514e-05 | 242481 | 1330.1496 | 7398379.8676 | 5798.78 | 825.7838 | 0.86914 | 0.0060898 | -0.0012765 | -0.021124 | -0.016311 | 5799.7258 | 825.6491 | 7397173.3406 | 0.86907 | 0.0060935 | -0.0012767 | -0.02112 | -0.016304 | 5799.7254 | 825.6492 | 13.4046
| 3 | 2023-11-09 08:43 | 182.5525 | 0.96676 | 19.9359 | 27.8421 | 13.4032 | 62.9039 | 1.5179e-05 | 273838 | 1500.051 | 111.8991 | 15.0005 | 11.9173 | 19.9359 | 28.8089 | 62.9061 | 241.1485 | 1.2698 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 15.0083 | 11.9752 | 13.4032 | [6 empty] | 15.0303 | 12.1399 | 13.4032 | [6 empty] | 20.2781 | 26.2189 | 825.8039 | 58.4393 | 1.5145e-05 | 242828 | 1330.182 | 7394916.6666 | 5798.78 | 825.8039 | 0.86893 | 0.0061005 | -0.0012766 | -0.021113 | -0.016289 | 5799.7245 | 825.6694 | 7393712.3337 | 0.86886 | 0.0061043 | -0.0012768 | -0.021109 | -0.016282 | 5799.7241 | 825.6695 | 13.4032
| 4 | 2023-11-09 08:46 | 182.2933 | 0.96673 | 19.968 | 27.8516 | 13.4039 | 62.9167 | 1.5181e-05 | 260396 | 1428.4453 | 106.5519 | 14.2845 | 6.5693 | 19.968 | 28.8183 | 62.9189 | 241.163 | 1.2698 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 14.2918 | 6.6244 | 13.4039 | [6 empty] | 14.3129 | 6.7813 | 13.4039 | [6 empty] | 20.3307 | 26.2284 | 825.7588 | 58.4456 | 1.5148e-05 | 242470 | 1330.1093 | 7393963.2687 | 5798.78 | 825.7588 | 0.86888 | 0.0061035 | -0.0012765 | -0.021106 | -0.016279 | 5799.724 | 825.6244 | 7392759.779 | 0.86881 | 0.0061073 | -0.0012768 | -0.021103 | -0.016272 | 5799.7236 | 825.6244 | 13.4039
| 6 | 2023-11-09 09:01 | 182.3177 | 0.9667 | 20.0542 | 27.7182 | 11.5438 | 62.5206 | 1.5181e-05 | 211490 | 1160.0079 | 100.4788 | 11.6001 | 0.4877 | 20.0542 | 28.6849 | 62.5227 | 241.37 | 1.2698 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 11.606 | 0.53894 | 11.5438 |  |  | 0.80965 |  |  |  | 11.623 | 0.68656 | 11.5438 |  |  | 0.81083 |  |  |  | 20.3666 | 26.541 | 701.394 | 59.2553 | 1.5159e-05 | 205980 | 1129.7859 | 6362747.0996 | 5798.78 | 701.394 | 0.80364 | 0.010564 | -0.0014902 | -0.017969 | -0.0088955 | 5799.2958 | 701.3316 | 6362181.1518 | 0.80361 | 0.010567 | -0.0014904 | -0.017968 | -0.0088909 | 5799.2956 | 701.3316 | 11.5438
| 7 | 2023-11-09 09:04 | 182.3013 | 0.96664 | 20.0168 | 27.7093 | 11.5431 | 62.5112 | 1.5179e-05 | 210842 | 1156.558 | 100.1857 | 11.5656 | 0.1946 | 20.0168 | 28.676 | 62.5133 | 241.3509 | 1.2698 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 11.5715 | 0.2456 | 11.5431 | [6 empty] | 11.5885 | 0.39281 | 11.5431 | [6 empty] | 20.3315 | 26.5322 | 701.4811 | 59.2445 | 1.5157e-05 | 205987 | 1129.9263 | 6363184.1928 | 5798.78 | 701.4811 | 0.80367 | 0.010562 | -0.0014903 | -0.017974 | -0.0089028 | 5799.2963 | 701.4187 | 6362617.7426 | 0.80364 | 0.010565 | -0.0014904 | -0.017973 | -0.0088981 | 5799.296 | 701.4187 | 11.5431
| 8 | 2023-11-09 09:08 | 182.6829 | 0.96661 | 20.0225 | 27.7064 | 11.5419 | 62.5007 | 1.5179e-05 | 214896 | 1176.3337 | 101.9099 | 11.7633 | 1.9189 | 20.0225 | 28.673 | 62.5029 | 241.3584 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 11.7693 | 1.9708 | 11.5419 | [6 empty] | 11.7866 | 2.1205 | 11.5419 | [6 empty] | 20.3216 | 26.5298 | 701.4364 | 59.2418 | 1.5156e-05 | 206405 | 1129.8542 | 6362715.8838 | 5798.78 | 701.4364 | 0.80364 | 0.010564 | -0.0014905 | -0.017974 | -0.0088999 | 5799.2961 | 701.3739 | 6362149.6558 | 0.8036 | 0.010567 | -0.0014906 | -0.017972 | -0.0088953 | 5799.2958 | 701.374 | 11.5419
| 9 | 2023-11-09 09:11 | 182.7577 | 0.96661 | 20.0006 | 27.7011 | 11.5397 | 62.495 | 1.5178e-05 | 211478 | 1157.1496 | 100.267 | 11.5715 | 0.27593 | 20.0006 | 28.6677 | 62.4972 | 241.3473 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99853 | 1.002 | 11.5774 | 0.32691 | 11.5397 | [6 empty] | 11.5944 | 0.47424 | 11.5397 | [6 empty] | 20.3105 | 26.5247 | 701.4141 | 59.2324 | 1.5156e-05 | 206483 | 1129.8183 | 6361781.4448 | 5798.78 | 701.4141 | 0.80358 | 0.01057 | -0.0014908 | -0.017974 | -0.0088949 | 5799.2958 | 701.3517 | 6361215.6233 | 0.80354 | 0.010573 | -0.0014909 | -0.017972 | -0.0088902 | 5799.2955 | 701.3517 | 11.5397
| 11 | 2023-11-09 09:26 | 183.5491 | 0.9666 | 20.027 | 27.5811 | 9.8279 | 62.1599 | 1.5175e-05 | 180792 | 984.9788 | 100.2226 | 9.8498 | 0.22242 | 20.027 | 28.5477 | 62.1621 | 241.4941 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 9.8547 | 0.27262 | 9.8279 | 9.8262 | 0.28259 | 0.10567 | 0.32152 | 0.25 | 0.40728 | 9.8692 | 0.41969 | 9.8279 | 9.8262 | 0.42965 | 0.10581 | 0.32195 | 0.25 | 0.40761 | 20.2126 | 26.7402 | 591.2686 | 59.8383 | 1.5158e-05 | 174812 | 952.399 | 5416916.5489 | 5798.78 | 591.2686 | 0.73375 | 0.017127 | -0.001762 | -0.015187 | 0.00017829 | 5798.7697 | 591.2696 | 5416926.2068 | 0.73375 | 0.017127 | -0.001762 | -0.015187 | 0.00017819 | 5798.7697 | 591.2696 | 9.8279
| 12 | 2023-11-09 09:30 | 183.4816 | 0.96659 | 20.0326 | 27.5765 | 9.826 | 62.1453 | 1.5175e-05 | 180527 | 983.8971 | 100.1323 | 9.839 | 0.13212 | 20.0326 | 28.543 | 62.1474 | 241.5032 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 9.8439 | 0.18225 | 9.826 | [6 empty] | 9.8583 | 0.32918 | 9.826 | [6 empty] | 20.2345 | 26.736 | 591.3407 | 59.8191 | 1.5159e-05 | 174769 | 952.5152 | 5415528.0495 | 5798.78 | 591.3407 | 0.73364 | 0.017139 | -0.0017623 | -0.015186 | 0.00019027 | 5798.769 | 591.3418 | 5415538.3537 | 0.73364 | 0.017139 | -0.0017623 | -0.015186 | 0.00019016 | 5798.769 | 591.3418 | 9.826
| 13 | 2023-11-09 09:33 | 183.6356 | 0.96654 | 20.0631 | 27.5768 | 9.8246 | 62.1342 | 1.5177e-05 | 181033 | 985.8271 | 100.3428 | 9.8583 | 0.34262 | 20.0631 | 28.5434 | 62.1363 | 241.5261 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 9.8632 | 0.39289 | 9.8246 | [6 empty] | 9.8777 | 0.54009 | 9.8246 | [6 empty] | 20.2323 | 26.7365 | 591.2368 | 59.8213 | 1.5159e-05 | 174885 | 952.3478 | 5414797.9652 | 5798.78 | 591.2368 | 0.73358 | 0.017145 | -0.0017626 | -0.015184 | 0.0001985 | 5798.7685 | 591.2379 | 5414808.7136 | 0.73358 | 0.017145 | -0.0017626 | -0.015184 | 0.00019838 | 5798.7685 | 591.2379 | 9.8246
| 15 | 2023-11-09 09:42 | 183.8474 | 0.96646 | 20.0706 | 27.4866 | 8.1042 | 61.8877 | 1.5174e-05 | 149137 | 811.2002 | 100.1081 | 8.112 | 0.096119 | 20.0706 | 28.453 | 61.8897 | 241.6268 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 8.116 | 0.14569 | 8.1042 | 8.1028 | 0.11725 | 0.10417 | 0.31696 | 0.25 | 0.40369 | 8.1279 | 0.29241 | 8.1042 | 8.1028 | 0.26398 | 0.10431 | 0.31737 | 0.25 | 0.40401 | 20.2597 | 26.9267 | 483.6457 | 60.3162 | 1.5166e-05 | 143225 | 779.043 | 4463989.0296 | 5798.78 | 483.6457 | 0.64972 | 0.026599 | -0.0021567 | -0.012438 | 0.012004 | 5798.0839 | 483.7038 | 4464524.966 | 0.64978 | 0.026593 | -0.0021564 | -0.01244 | 0.011997 | 5798.0843 | 483.7037 | 8.1042
| 16 | 2023-11-09 09:45 | 184.6397 | 0.96645 | 19.9762 | 27.4671 | 8.1024 | 61.8718 | 1.5169e-05 | 149835 | 811.4994 | 100.1669 | 8.115 | 0.15487 | 19.9762 | 28.4336 | 61.8739 | 241.5757 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 8.119 | 0.20426 | 8.1024 | [6 empty] | 8.1309 | 0.35113 | 8.1024 | [6 empty] | 20.1381 | 26.9076 | 483.5842 | 60.3107 | 1.5159e-05 | 143824 | 778.9441 | 4464857.6634 | 5798.78 | 483.5842 | 0.64981 | 0.026589 | -0.0021572 | -0.012444 | 0.011988 | 5798.0848 | 483.6422 | 4465392.973 | 0.64986 | 0.026583 | -0.0021569 | -0.012446 | 0.011981 | 5798.0853 | 483.6422 | 8.1024
| 17 | 2023-11-09 09:49 | 184.8076 | 0.9664 | 19.9438 | 27.4569 | 8.1018 | 61.8566 | 1.5167e-05 | 149656 | 809.7936 | 99.9646 | 8.0979 | -0.047372 | 19.9438 | 28.4233 | 61.8588 | 241.562 | 1.2699 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 8.1019 | 0.0018185 | 8.1018 | [6 empty] | 8.1138 | 0.14841 | 8.1018 | [6 empty] | 20.079 | 26.8971 | 483.5951 | 60.3044 | 1.5156e-05 | 143958 | 778.9616 | 4465397.0828 | 5798.78 | 483.5951 | 0.64986 | 0.026583 | -0.0021573 | -0.012448 | 0.011978 | 5798.0854 | 483.653 | 4465932.0037 | 0.64991 | 0.026577 | -0.0021571 | -0.012449 | 0.01197 | 5798.0859 | 483.653 | 8.1018
| 22 | 2023-11-09 10:12 | 185.124 | 0.96645 | 19.9283 | 27.3636 | 6.5441 | 61.6112 | 1.5164e-05 | 121357 | 655.5443 | 100.1994 | 6.5554 | 0.17409 | 19.9283 | 28.33 | 61.6132 | 241.6486 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 6.5586 | 0.22277 | 6.5441 | 6.5438 | 0.20947 | 0.12035 | 0.36619 | 0.25 | 0.44339 | 6.5682 | 0.36958 | 6.5441 | 6.5438 | 0.35626 | 0.12053 | 0.36673 | 0.25 | 0.44383 | 19.9442 | 27.0031 | 388.4134 | 60.6381 | 1.5153e-05 | 115822 | 625.6454 | 3607073.1222 | 5798.78 | 388.4134 | 0.55715 | 0.037959 | -0.002704 | -0.010015 | 0.02524 | 5797.3164 | 388.5114 | 3607983.7817 | 0.55726 | 0.037946 | -0.0027033 | -0.010018 | 0.025225 | 5797.3173 | 388.5114 | 6.5441
| 23 | 2023-11-09 10:18 | 186.2237 | 0.96638 | 19.9168 | 27.3566 | 6.5433 | 61.5968 | 1.5163e-05 | 121894 | 654.5568 | 100.0597 | 6.5456 | 0.034466 | 19.9168 | 28.323 | 61.5987 | 241.6473 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 6.5487 | 0.083026 | 6.5433 | [6 empty] | 6.5583 | 0.22963 | 6.5433 | [6 empty] | 19.936 | 26.9957 | 388.48 | 60.6208 | 1.5153e-05 | 116530 | 625.7527 | 3606812.9977 | 5798.78 | 388.48 | 0.55712 | 0.037963 | -0.0027043 | -0.010017 | 0.025242 | 5797.3163 | 388.578 | 3607723.665 | 0.55723 | 0.037949 | -0.0027036 | -0.010019 | 0.025227 | 5797.3172 | 388.578 | 6.5433
| 24 | 2023-11-09 10:22 | 181.7465 | 0.96635 | 19.9096 | 27.3529 | 6.544 | 61.5894 | 1.5163e-05 | 119261 | 656.1941 | 100.2993 | 6.5619 | 0.27398 | 19.9096 | 28.3192 | 61.5915 | 241.6458 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.002 | 6.5651 | 0.32262 | 6.544 | [6 empty] | 6.5747 | 0.46958 | 6.544 | [6 empty] | 19.9302 | 26.9922 | 388.5675 | 60.6137 | 1.5152e-05 | 113754 | 625.8937 | 3607289.1712 | 5798.78 | 388.5675 | 0.55718 | 0.037956 | -0.002704 | -0.010019 | 0.025233 | 5797.3168 | 388.6655 | 3608199.6301 | 0.55729 | 0.037942 | -0.0027032 | -0.010021 | 0.025218 | 5797.3177 | 388.6655 | 6.544
| 26 | 2023-11-09 10:33 | 186.9056 | 0.9663 | 20.0112 | 27.3076 | 4.7566 | 61.4283 | 1.5166e-05 | 88871 | 475.4861 | 100.0064 | 4.7549 | -0.037137 | 20.0112 | 28.2739 | 61.4302 | 241.7703 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 4.7572 | 0.011158 | 4.7566 | 4.756 | -0.020361 | 0.098543 | 0.29984 | 0.25 | 0.39039 | 4.7641 | 0.15751 | 4.7566 | 4.756 | 0.12595 | 0.098688 | 0.30028 | 0.25 | 0.39072 | 19.9743 | 27.1236 | 280.8322 | 60.9489 | 1.5158e-05 | 84548 | 452.3568 | 2620479.0703 | 5798.78 | 280.8322 | 0.41838 | 0.054585 | -0.0038086 | -0.007247 | 0.043529 | 5796.2559 | 280.9545 | 2621620.2322 | 0.41857 | 0.054563 | -0.0038068 | -0.0072502 | 0.043506 | 5796.2572 | 280.9545 | 4.7566
| 27 | 2023-11-09 10:36 | 187.4586 | 0.96628 | 19.9995 | 27.3012 | 4.7557 | 61.4154 | 1.5165e-05 | 89158 | 475.6142 | 100.0538 | 4.7561 | 0.010302 | 19.9995 | 28.2674 | 61.4175 | 241.7682 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 4.7584 | 0.058568 | 4.7557 | [6 empty] | 4.7654 | 0.20499 | 4.7557 | [6 empty] | 19.9469 | 27.1171 | 280.8084 | 60.9415 | 1.5157e-05 | 84791 | 452.3184 | 2620198.2607 | 5798.78 | 280.8084 | 0.41833 | 0.05459 | -0.0038094 | -0.0072472 | 0.043533 | 5796.2556 | 280.9307 | 2621339.4109 | 0.41852 | 0.054569 | -0.0038076 | -0.0072504 | 0.043511 | 5796.2569 | 280.9307 | 4.7557
| 28 | 2023-11-09 10:40 | 181.7978 | 0.96623 | 19.9948 | 27.2963 | 4.7558 | 61.4038 | 1.5165e-05 | 86304 | 474.7251 | 99.8645 | 4.7473 | -0.17895 | 19.9948 | 28.2625 | 61.406 | 241.7699 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 4.7495 | -0.13081 | 4.7558 | [6 empty] | 4.7565 | 0.015338 | 4.7558 | [6 empty] | 19.9222 | 27.1124 | 280.8306 | 60.9381 | 1.5156e-05 | 82237 | 452.3541 | 2620482.7293 | 5798.78 | 280.8306 | 0.41838 | 0.054584 | -0.0038093 | -0.0072486 | 0.043527 | 5796.256 | 280.9529 | 2621623.8308 | 0.41857 | 0.054563 | -0.0038075 | -0.0072518 | 0.043504 | 5796.2573 | 280.9528 | 4.7558
| 30 | 2023-11-09 10:49 | 187.3821 | 0.96624 | 20.0448 | 27.2647 | 3.2796 | 61.3 | 1.5166e-05 | 61472 | 328.057 | 100.0896 | 3.2806 | 0.03006 | 20.0448 | 28.2309 | 61.302 | 241.8408 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 3.2821 | 0.078142 | 3.2796 | 3.2797 | 0.084887 | 0.017753 | 0.054017 | 0.25 | 0.25577 | 3.2869 | 0.22451 | 3.2796 | 3.2797 | 0.23129 | 0.017775 | 0.054084 | 0.25 | 0.25578 | 19.9862 | 27.184 | 193.0989 | 61.1059 | 1.5161e-05 | 58283 | 311.0383 | 1806180.1363 | 5798.78 | 193.0989 | 0.25676 | 0.070253 | -0.0057338 | -0.0049852 | 0.059534 | 5795.3277 | 193.2139 | 1807256.0724 | 0.25702 | 0.070233 | -0.00573 | -0.0049881 | 0.059515 | 5795.3289 | 193.2139 | 3.2796
| 31 | 2023-11-09 10:53 | 186.6416 | 0.96625 | 19.9964 | 27.2555 | 3.2796 | 61.2939 | 1.5164e-05 | 61246 | 328.1476 | 100.1166 | 3.2815 | 0.057033 | 19.9964 | 28.2218 | 61.296 | 241.8139 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 3.283 | 0.10502 | 3.2796 | [6 empty] | 3.2879 | 0.25146 | 3.2796 | [6 empty] | 19.9205 | 27.1747 | 193.0999 | 61.106 | 1.5158e-05 | 58053 | 311.0399 | 1806590.9084 | 5798.78 | 193.0999 | 0.25686 | 0.070245 | -0.0057338 | -0.0049868 | 0.059525 | 5795.3283 | 193.2149 | 1807666.9183 | 0.25712 | 0.070225 | -0.00573 | -0.0049898 | 0.059505 | 5795.3294 | 193.2149 | 3.2796
| 32 | 2023-11-09 10:56 | 187.2341 | 0.96627 | 19.9563 | 27.248 | 3.2799 | 61.2893 | 1.5161e-05 | 61426 | 328.0706 | 100.0831 | 3.2807 | 0.023605 | 19.9563 | 28.2143 | 61.2911 | 241.7915 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 3.2823 | 0.071495 | 3.2799 | [6 empty] | 3.2871 | 0.21791 | 3.2799 | [6 empty] | 19.857 | 27.1673 | 193.1056 | 61.1102 | 1.5154e-05 | 58239 | 311.0491 | 1807150.6527 | 5798.78 | 193.1056 | 0.25699 | 0.070235 | -0.0057332 | -0.0049886 | 0.059513 | 5795.329 | 193.2206 | 1808226.7811 | 0.25725 | 0.070214 | -0.0057294 | -0.0049916 | 0.059493 | 5795.3301 | 193.2206 | 3.2799
| 34 | 2023-11-09 11:04 | 181.8117 | 0.96622 | 20.0133 | 27.2321 | 1.5837 | 61.2245 | 1.5164e-05 | 28820 | 158.5156 | 100.1611 | 1.5852 | 0.094312 | 20.0133 | 28.1983 | 61.2265 | 241.8512 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 1.5859 | 0.14219 | 1.5837 | 1.5833 | 0.16111 | 0.023145 | 0.070424 | 0.25 | 0.25973 | 1.5882 | 0.28863 | 1.5837 | 1.5833 | 0.3076 | 0.023211 | 0.070624 | 0.25 | 0.25978 | 19.8229 | 27.2238 | 92.9804 | 61.2751 | 1.5155e-05 | 27230 | 149.7703 | 872480.8371 | 5798.78 | 92.9804 | -0.059244 | 0.082394 | -0.013367 | -0.0024037 | 0.066623 | 5794.9167 | 93.0424 | 873062.4983 | -0.058955 | 0.082394 | -0.013356 | -0.0024053 | 0.066632 | 5794.9161 | 93.0424 | 1.5837
| 35 | 2023-11-09 11:09 | 190.3713 | 0.96623 | 19.9259 | 27.2165 | 1.5831 | 61.2162 | 1.5159e-05 | 30180 | 158.5323 | 100.206 | 1.5853 | 0.1392 | 19.9259 | 28.1828 | 61.2182 | 241.8017 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 1.5861 | 0.18691 | 1.5831 | [6 empty] | 1.5884 | 0.33348 | 1.5831 | [6 empty] | 19.8053 | 27.2087 | 93.0001 | 61.2411 | 1.5153e-05 | 28518 | 149.802 | 872255.9613 | 5798.78 | 93.0001 | -0.059356 | 0.082393 | -0.013371 | -0.0024042 | 0.066619 | 5794.9169 | 93.0621 | 872837.4334 | -0.059067 | 0.082394 | -0.01336 | -0.0024058 | 0.066628 | 5794.9164 | 93.0621 | 1.5831
| 36 | 2023-11-09 11:13 | 188.193 | 0.96623 | 19.98 | 27.2176 | 1.5831 | 61.1983 | 1.5162e-05 | 29825 | 158.4809 | 100.1732 | 1.5848 | 0.10647 | 19.98 | 28.1839 | 61.2004 | 241.8413 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 1.5856 | 0.15422 | 1.5831 | [6 empty] | 1.5879 | 0.3007 | 1.5831 | [6 empty] | 19.8006 | 27.2095 | 92.9946 | 61.245 | 1.5153e-05 | 28190 | 149.7931 | 872269.6194 | 5798.78 | 92.9946 | -0.059349 | 0.082393 | -0.013371 | -0.0024041 | 0.066619 | 5794.9169 | 93.0566 | 872851.0999 | -0.05906 | 0.082394 | -0.01336 | -0.0024057 | 0.066628 | 5794.9164 | 93.0566 | 1.5831
| 37 | 2023-11-09 11:28 | 61.8186 | 0.96602 | 19.7757 | 27.208 | #NUM! | 61.2501 | 1.5152e-05 | 8 | 0.12941 | NaN | 0.0012941 | #NUM! | 19.7757 | 28.174 | 61.252 | 241.6976 | 1.27 | 93.14 | 44.7 | 0.02235 | 1 | 0.99854 | 1.0019 | 0.0012947 | #NUM! | #NUM! | [6 empty] | 0.0012966 | #NUM! | #NUM! | [6 empty] | 19.6667 | 27.2227 | 0 | 61.3312 | 1.5147e-05 | 0 | 0 | 0 | 5798.78 | 0 | #NUM! | #NUM! | #DIV/0! | 0 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM!
0.12316

## (TBL) Coriolis CO2 37 bar
Table: Coriolis | [57 empty] | a0 | 0.08203
[50 empty] | Fitted K-factor |  | GC Origional | [5 empty] | a1 | -0.012922
Job Number | EMR013 | [47 empty] | 0.1524 | 0 | 9872303.3667 | -4e-07 | [5 empty] | a2 | -0.11908
Test Number | T03135 | [48 empty] | 5798.78 | 833097.815 | 5802.3 | [5 empty] | a3 | -0.067147
Test Title | 3" Coriolis Intercomparison 37 Bar CO2 | [25 empty] | 0 | [30 empty] | a4 | 0.1215
Test Description | 3" Coriolis Intercomparison 37 Bar CO2 | [56 empty] | b0 | -0.00011761
[39 empty] | 0.28073 | [18 empty] | b1 | -0.016601
[58 empty] | cp | -0.042815
ONLY PRESSURE CORRECTION APPLIED - NO COMPRESSIBILITY
[15 empty] | RefProp version 10 | [13 empty] | Averaged values | [8 empty] | Averaged values | [5 empty] | Reference Turbine Calcs
| Test Point | TestPoint Timestamp | Collection Time (t) | Ambient Pressure (Pamb) | Test Temperature (T) | Test Pressure (Pgauge) | Reference Mass Flow Rate (ṁ) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | K Factor (K) | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Test Temperature (T) | Test Pressure (Pabs) | Gas Density (ρ) | Speed of Sound | Isentropic expansion | Tube Frequency | Tube Diameter | Tube Radius | a | Compressibility correction fc | Pressure correction fp | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Meter Est Mass Flow (ṁi) | Meter % Error (E) | Reference Mass Flow Rate (ṁ) | Reference Mass Flow Rate (ṁ) | Meter % Error (E) | Standard Dev | Uncertainty of mean | Uncertainty of turbine mass | Overall Uncertainty | Test Temperature (T) | Test Pressure (Pgauge) | Reference Volume Flow Rate (Q) | Test Gas Density (ρ) | Test Gas Dynamic Viscosity (μ) | Meter Signal (P) | Test Meter Frequency (f) | Reynolds | K nominal | Meter Nominal Flow (Q) | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Meter Est Volume Flow (Qi) | Reynolds | LOG(Re) | eRE | eB | eP | eTM | K Factor (K) | Reference Volume Flow Rate (Q) | Reference Mass Flow Rate (ṁ)
| - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | % | °C | bara | kg/m³ | m/s | - | Hz | mm | mm | - | - | - | kg/s | % | kg/s | kg/s | % | % | % | % | % | kg/s | % | kg/s | kg/s | % | % | % | % | % | °C | barg | m³/h | kg/m³ | Pa s | pulse | Hz | - | - | m3/s | - | % | % | % | % | pulse/m³ | m³/h | - | - | % | % | % | % | pulse/m³ | m³/h | kg/s
| 3 | 2023-11-09 14:27 | 182.3444 | 0.96491 | 20.0088 | 37.0183 | 18.2673 | 90.4994 | 1.5548e-05 | 364829 | 2000.7691 | 109.5071 | 20.0077 | 9.5274 | 20.0088 | 37.9832 | 90.5006 | 230.7561 | 1.2687 | 92.65 | 44.7 | 0.02235 | 1 | 0.99841 | 1.0027 | 20.0298 | 9.6487 | 18.2673 |  |  | 1.4537 |  |  |  | 20.0617 | 9.823 | 18.2673 |  |  | 1.456 |  |  |  | 20.2696 | 34.9344 | 787.637 | 83.5086 | 1.5462e-05 | 231341 | 1268.7038 | 9872303.3667 | 5798.78 | 787.637 | 0.99442 | 0.0042035 | -0.00093802 | -0.021791 | -0.018525 | 5799.8542 | 787.4911 | 9870474.8493 | 0.99434 | 0.0042012 | -0.0009382 | -0.021787 | -0.018524 | 5799.8541 | 787.4911 | 18.2673
| 4 | 2023-11-09 14:31 | 181.767 | 0.96484 | 20.0402 | 37.031 | 18.2582 | 90.5178 | 1.555e-05 | 365002 | 2008.0762 | 109.9618 | 20.0808 | 9.9822 | 20.0402 | 37.9958 | 90.5188 | 230.7718 | 1.2687 | 92.65 | 44.7 | 0.02235 | 1 | 0.99841 | 1.0027 | 20.103 | 10.1041 | 18.2582 | [6 empty] | 20.135 | 10.2791 | 18.2582 | [6 empty] | 20.4141 | 34.9462 | 787.7856 | 83.4512 | 1.5468e-05 | 230652 | 1268.9431 | 9863199.9601 | 5798.78 | 787.7856 | 0.99402 | 0.004192 | -0.00093848 | -0.021773 | -0.018519 | 5799.8539 | 787.6397 | 9861373.7144 | 0.99394 | 0.0041897 | -0.00093866 | -0.021769 | -0.018518 | 5799.8538 | 787.6397 | 18.2582
| 5 | 2023-11-09 14:35 | 181.7742 | 0.96475 | 20.0467 | 37.0373 | 18.2529 | 90.5341 | 1.5551e-05 | 355914 | 1958.0007 | 107.2507 | 19.58 | 7.2706 | 20.0467 | 38.0021 | 90.5348 | 230.7706 | 1.2687 | 92.65 | 44.7 | 0.02235 | 1 | 0.99841 | 1.0027 | 19.6017 | 7.3895 | 18.2529 | [6 empty] | 19.6329 | 7.5602 | 18.2529 | [6 empty] | 20.4917 | 34.9513 | 787.8843 | 83.4166 | 1.5472e-05 | 230690 | 1269.1021 | 9858139.9076 | 5798.78 | 787.8843 | 0.99379 | 0.0041856 | -0.00093875 | -0.021763 | -0.018516 | 5799.8537 | 787.7384 | 9856314.8666 | 0.99371 | 0.0041833 | -0.00093893 | -0.021759 | -0.018515 | 5799.8536 | 787.7384 | 18.2529
| 7 | 2023-11-09 14:55 | 181.7688 | 0.96479 | 19.9369 | 36.6905 | 14.7071 | 89.4551 | 1.5529e-05 | 269667 | 1483.5715 | 100.8605 | 14.8357 | 0.87438 | 19.9369 | 37.6553 | 89.4565 | 231.0886 | 1.2687 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 14.8518 | 0.98396 | 14.7071 |  |  | 0.21622 |  |  |  | 14.8754 | 1.144 | 14.7071 |  |  | 0.21655 |  |  |  | 20.3679 | 35.3729 | 624.1476 | 84.8404 | 1.5485e-05 | 182743 | 1005.3596 | 7935754.7748 | 5798.78 | 624.1476 | 0.89959 | 0.0047236 | -0.0011747 | -0.017331 | -0.013782 | 5799.5792 | 624.0616 | 7934661.1837 | 0.89953 | 0.0047257 | -0.0011749 | -0.017329 | -0.013778 | 5799.579 | 624.0616 | 14.7071
| 8 | 2023-11-09 14:59 | 182.2481 | 0.96483 | 19.9005 | 36.6749 | 14.7058 | 89.4304 | 1.5527e-05 | 269225 | 1477.2449 | 100.4391 | 14.7724 | 0.4529 | 19.9005 | 37.6397 | 89.4317 | 231.0719 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 14.7885 | 0.56189 | 14.7058 | [6 empty] | 14.8119 | 0.7213 | 14.7058 | [6 empty] | 20.2926 | 35.3583 | 624.063 | 84.8445 | 1.5482e-05 | 183200 | 1005.2234 | 7937002.4063 | 5798.78 | 624.063 | 0.89966 | 0.0047211 | -0.0011748 | -0.017337 | -0.01379 | 5799.5797 | 623.977 | 7935908.006 | 0.8996 | 0.0047233 | -0.001175 | -0.017334 | -0.013786 | 5799.5794 | 623.977 | 14.7058
| 9 | 2023-11-09 15:03 | 181.7548 | 0.96488 | 19.9071 | 36.6722 | 14.7081 | 89.4166 | 1.5527e-05 | 268883 | 1479.372 | 100.5684 | 14.7937 | 0.58232 | 19.9071 | 37.6371 | 89.418 | 231.0816 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 14.8098 | 0.69143 | 14.7081 | [6 empty] | 14.8332 | 0.85103 | 14.7081 | [6 empty] | 20.2713 | 35.3555 | 624.117 | 84.8501 | 1.5481e-05 | 182720 | 1005.3103 | 7938721.7162 | 5798.78 | 624.117 | 0.89975 | 0.0047177 | -0.0011746 | -0.017341 | -0.013798 | 5799.5801 | 624.0309 | 7937626.5101 | 0.89969 | 0.0047198 | -0.0011748 | -0.017338 | -0.013793 | 5799.5798 | 624.0309 | 14.7081
| 10 | 2023-11-09 15:18 | 182.3011 | 0.96494 | 19.8984 | 36.4831 | 12.229 | 88.7935 | 1.5518e-05 | 223301 | 1224.9024 | 100.1576 | 12.249 | 0.16408 | 19.8984 | 37.448 | 88.795 | 231.3051 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 12.2622 | 0.27158 | 12.229 | 12.2268 | 0.19392 | 0.068427 | 0.2082 | 0.25 | 0.32534 | 12.2816 | 0.43021 | 12.229 | 12.2268 | 0.35243 | 0.068539 | 0.20854 | 0.25 | 0.32556 | 20.1853 | 35.5874 | 514.0042 | 85.6551 | 1.5488e-05 | 150935 | 827.9436 | 6597158.7689 | 5798.78 | 514.0042 | 0.81936 | 0.0093211 | -0.0014248 | -0.014325 | -0.0064289 | 5799.1528 | 513.9711 | 6596734.6717 | 0.81933 | 0.0093232 | -0.0014249 | -0.014324 | -0.0064259 | 5799.1526 | 513.9711 | 12.229
| 11 | 2023-11-09 15:22 | 181.8103 | 0.96495 | 19.8949 | 36.4747 | 12.2264 | 88.7682 | 1.5517e-05 | 222366 | 1223.066 | 100.0287 | 12.2307 | 0.03517 | 19.8949 | 37.4396 | 88.7697 | 231.3119 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 12.2438 | 0.14248 | 12.2264 | [6 empty] | 12.2631 | 0.3009 | 12.2264 | [6 empty] | 20.1461 | 35.5792 | 513.8926 | 85.6555 | 1.5486e-05 | 150496 | 827.7639 | 6596603.0551 | 5798.78 | 513.8926 | 0.81932 | 0.0093239 | -0.0014251 | -0.014325 | -0.0064266 | 5799.1527 | 513.8596 | 6596179.1444 | 0.81929 | 0.009326 | -0.0014252 | -0.014325 | -0.0064237 | 5799.1525 | 513.8596 | 12.2264
| 12 | 2023-11-09 15:26 | 181.811 | 0.96494 | 19.8777 | 36.4658 | 12.2251 | 88.7513 | 1.5516e-05 | 222401 | 1223.2535 | 100.054 | 12.2325 | 0.060448 | 19.8777 | 37.4307 | 88.753 | 231.3059 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0027 | 12.2456 | 0.16771 | 12.2251 | [6 empty] | 12.265 | 0.32617 | 12.2251 | [6 empty] | 20.122 | 35.5706 | 513.9075 | 85.6445 | 1.5484e-05 | 150501 | 827.788 | 6596537.608 | 5798.78 | 513.9075 | 0.81932 | 0.0093242 | -0.0014252 | -0.014327 | -0.0064284 | 5799.1528 | 513.8745 | 6596113.5847 | 0.81929 | 0.0093263 | -0.0014253 | -0.014326 | -0.0064254 | 5799.1526 | 513.8745 | 12.2251
| 13 | 2023-11-09 15:50 | 181.7627 | 0.96487 | 20.0105 | 36.3912 | 10.9176 | 88.4061 | 1.5518e-05 | 198087 | 1089.811 | 99.8207 | 10.8981 | -0.17846 | 20.0105 | 37.3561 | 88.4076 | 231.5272 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0026 | 10.9098 | -0.071733 | 10.9176 | 10.916 | 0.049377 | 0.10528 | 0.32033 | 0.25 | 0.40634 | 10.927 | 0.08605 | 10.9176 | 10.916 | 0.20734 | 0.10544 | 0.32082 | 0.25 | 0.40672 | 20.1614 | 35.6845 | 457.0967 | 85.9854 | 1.5491e-05 | 133828 | 736.2786 | 5888071.9044 | 5798.78 | 457.0967 | 0.76997 | 0.013534 | -0.0016054 | -0.012755 | -0.00082575 | 5798.8279 | 457.0929 | 5888023.2842 | 0.76997 | 0.013534 | -0.0016054 | -0.012754 | -0.00082532 | 5798.8279 | 457.0929 | 10.9176
| 14 | 2023-11-09 15:53 | 181.7672 | 0.96493 | 20.0168 | 36.3901 | 10.9161 | 88.3981 | 1.5518e-05 | 198443 | 1091.7428 | 100.0113 | 10.9174 | 0.012118 | 20.0168 | 37.3551 | 88.3997 | 231.5345 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0026 | 10.9291 | 0.11905 | 10.9161 | [6 empty] | 10.9464 | 0.27713 | 10.9161 | [6 empty] | 20.1769 | 35.6836 | 457.106 | 85.972 | 1.5492e-05 | 133834 | 736.2936 | 5887040.7224 | 5798.78 | 457.106 | 0.7699 | 0.013541 | -0.0016056 | -0.012753 | -0.00081738 | 5798.8274 | 457.1022 | 5886992.6033 | 0.76989 | 0.013542 | -0.0016056 | -0.012753 | -0.00081696 | 5798.8274 | 457.1022 | 10.9161
| 15 | 2023-11-09 15:57 | 181.7622 | 0.96496 | 20.0185 | 36.3886 | 10.9142 | 88.392 | 1.5518e-05 | 198367 | 1091.3547 | 99.9931 | 10.9135 | -0.0060978 | 20.0185 | 37.3536 | 88.3934 | 231.538 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99842 | 1.0026 | 10.9252 | 0.10081 | 10.9142 | [6 empty] | 10.9425 | 0.25885 | 10.9142 | [6 empty] | 20.1779 | 35.6817 | 457.0639 | 85.965 | 1.5492e-05 | 133818 | 736.2258 | 5886039.1902 | 5798.78 | 457.0639 | 0.76982 | 0.013548 | -0.0016059 | -0.012752 | -0.00080918 | 5798.8269 | 457.0602 | 5885991.562 | 0.76982 | 0.013549 | -0.0016059 | -0.012751 | -0.00080876 | 5798.8269 | 457.0602 | 10.9142
| 16 | 2023-11-09 16:13 | 182.1892 | 0.96505 | 20.0888 | 36.3197 | 9.3987 | 88.1128 | 1.5518e-05 | 171163 | 939.4792 | 99.9657 | 9.3948 | -0.041674 | 20.0888 | 37.2847 | 88.1148 | 231.6902 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 9.4048 | 0.064872 | 9.3987 | 9.3974 | 0.13938 | 0.095191 | 0.18685 | 0.25 | 0.31211 | 9.4196 | 0.22265 | 9.3987 | 9.3974 | 0.29727 | 0.095338 | 0.18714 | 0.25 | 0.31228 | 20.2237 | 35.7995 | 391.9679 | 86.3153 | 1.5499e-05 | 115029 | 631.371 | 5065923.2142 | 5798.78 | 391.9679 | 0.70466 | 0.020257 | -0.0018813 | -0.010945 | 0.0074304 | 5798.3491 | 391.997 | 5066299.6595 | 0.70469 | 0.020253 | -0.0018812 | -0.010946 | 0.0074261 | 5798.3494 | 391.997 | 9.3987
| 17 | 2023-11-09 16:17 | 124.8905 | 0.96505 | 20.0854 | 36.3146 | 9.3979 | 88.0988 | 1.5517e-05 | 117413 | 940.1279 | 100.0429 | 9.4013 | 0.035507 | 20.0854 | 37.2797 | 88.1007 | 231.6931 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 9.4113 | 0.1421 | 9.3979 | [6 empty] | 9.4261 | 0.29999 | 9.3979 | [6 empty] | 20.2041 | 35.7949 | 391.942 | 86.314 | 1.5498e-05 | 78847 | 631.3292 | 5065841.4876 | 5798.78 | 391.942 | 0.70465 | 0.020257 | -0.0018815 | -0.010945 | 0.0074305 | 5798.3491 | 391.9711 | 5066217.933 | 0.70468 | 0.020254 | -0.0018814 | -0.010946 | 0.0074263 | 5798.3494 | 391.9711 | 9.3979
| 18 | 2023-11-09 16:21 | 181.769 | 0.96504 | 20.0801 | 36.3107 | 9.3961 | 88.0902 | 1.5517e-05 | 170826 | 939.797 | 100.0276 | 9.398 | 0.020182 | 20.0801 | 37.2758 | 88.0916 | 231.6927 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 9.408 | 0.12673 | 9.3961 | [6 empty] | 9.4228 | 0.2846 | 9.3961 | [6 empty] | 20.197 | 35.791 | 391.9006 | 86.3059 | 1.5497e-05 | 114744 | 631.2626 | 5064989.0765 | 5798.78 | 391.9006 | 0.70458 | 0.020265 | -0.0018819 | -0.010944 | 0.0074392 | 5798.3486 | 391.9298 | 5065365.8997 | 0.70461 | 0.020262 | -0.0018817 | -0.010945 | 0.007435 | 5798.3489 | 391.9298 | 9.3961
| 19 | 2023-11-09 16:24 | 181.7819 | 0.96508 | 20.0748 | 36.3072 | 9.397 | 88.082 | 1.5517e-05 | 170751 | 939.3177 | 99.9665 | 9.3932 | -0.040944 | 20.0748 | 37.2723 | 88.084 | 231.6918 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 9.4032 | 0.065509 | 9.397 | [6 empty] | 9.418 | 0.22328 | 9.397 | [6 empty] | 20.1899 | 35.7875 | 391.9684 | 86.2997 | 1.5497e-05 | 114772 | 631.3719 | 5065648.8505 | 5798.78 | 391.9684 | 0.70464 | 0.020259 | -0.0018817 | -0.010946 | 0.007431 | 5798.3491 | 391.9976 | 5066025.3078 | 0.70467 | 0.020256 | -0.0018815 | -0.010947 | 0.0074268 | 5798.3493 | 391.9975 | 9.397
| 20 | 2023-11-09 16:28 | 181.7931 | 0.96513 | 20.0771 | 36.3051 | 9.3974 | 88.0734 | 1.5517e-05 | 171164 | 941.5318 | 100.1985 | 9.4153 | 0.19102 | 20.0771 | 37.2702 | 88.0754 | 231.6966 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 9.4253 | 0.29771 | 9.3974 | [6 empty] | 9.4402 | 0.45585 | 9.3974 | [6 empty] | 20.1887 | 35.7855 | 392.0058 | 86.2947 | 1.5497e-05 | 114790 | 631.432 | 5065877.1033 | 5798.78 | 392.0058 | 0.70465 | 0.020257 | -0.0018816 | -0.010947 | 0.007428 | 5798.3493 | 392.0349 | 5066253.4224 | 0.70469 | 0.020253 | -0.0018815 | -0.010948 | 0.0074237 | 5798.3495 | 392.0349 | 9.3974
| 21 | 2023-11-09 16:43 | 183.2483 | 0.9652 | 19.9882 | 36.2061 | 7.6467 | 87.8127 | 1.5508e-05 | 140118 | 764.6347 | 100.0152 | 7.6463 | -0.0043853 | 19.9882 | 37.1713 | 87.8142 | 231.7303 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 7.6544 | 0.10138 | 7.6467 | 7.6468 | 0.0998 | 0.10858 | 0.33038 | 0.25 | 0.41431 | 7.6665 | 0.25916 | 7.6467 | 7.6468 | 0.2576 | 0.10874 | 0.33087 | 0.25 | 0.41469 | 20.0904 | 35.8669 | 317.7105 | 86.6282 | 1.5497e-05 | 93779 | 511.7592 | 4121699.649 | 5798.78 | 317.7105 | 0.61508 | 0.030795 | -0.0023457 | -0.0088872 | 0.019563 | 5797.6456 | 317.7727 | 4122506.1201 | 0.61516 | 0.030785 | -0.0023452 | -0.0088889 | 0.019551 | 5797.6463 | 317.7726 | 7.6467
| 22 | 2023-11-09 16:47 | 181.7883 | 0.96522 | 19.9578 | 36.1955 | 7.6474 | 87.7997 | 1.5507e-05 | 139162 | 765.5167 | 100.1214 | 7.6552 | 0.10182 | 19.9578 | 37.1607 | 87.8013 | 231.7138 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 7.6633 | 0.20759 | 7.6474 | [6 empty] | 7.6753 | 0.36555 | 7.6474 | [6 empty] | 20.0483 | 35.8564 | 317.7554 | 86.6238 | 1.5494e-05 | 93045 | 511.8315 | 4122663.5032 | 5798.78 | 317.7554 | 0.61518 | 0.030783 | -0.0023455 | -0.0088905 | 0.019547 | 5797.6465 | 317.8175 | 4123469.5217 | 0.61526 | 0.030773 | -0.002345 | -0.0088922 | 0.019535 | 5797.6472 | 317.8175 | 7.6474
| 23 | 2023-11-09 16:51 | 181.8245 | 0.96521 | 19.929 | 36.1865 | 7.6462 | 87.7908 | 1.5505e-05 | 138867 | 763.7417 | 99.9045 | 7.6374 | -0.11501 | 19.929 | 37.1517 | 87.7927 | 231.6966 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 7.6455 | -0.0095616 | 7.6462 | [6 empty] | 7.6575 | 0.14809 | 7.6462 | [6 empty] | 20.0208 | 35.8476 | 317.7399 | 86.6148 | 1.5493e-05 | 93059 | 511.8066 | 4122439.9091 | 5798.78 | 317.7399 | 0.61515 | 0.030786 | -0.0023458 | -0.0088913 | 0.019549 | 5797.6464 | 317.802 | 4123245.9564 | 0.61524 | 0.030776 | -0.0023453 | -0.008893 | 0.019537 | 5797.6471 | 317.802 | 7.6462
| 24 | 2023-11-09 16:57 | 181.7996 | 0.9653 | 19.9856 | 36.1496 | 6.2561 | 87.6289 | 1.5506e-05 | 113730 | 625.5791 | 100.0263 | 6.2558 | -0.0050875 | 19.9856 | 37.1149 | 87.6307 | 231.7963 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 6.2624 | 0.10033 | 6.2561 | 6.2563 | 0.11258 | 0.015465 | 0.047054 | 0.25 | 0.25439 | 6.2723 | 0.25801 | 6.2561 | 6.2563 | 0.2703 | 0.015494 | 0.047142 | 0.25 | 0.25441 | 20.0334 | 35.9273 | 259.1914 | 86.866 | 1.5497e-05 | 75901 | 417.4983 | 3371650.4647 | 5798.78 | 259.1914 | 0.52784 | 0.041587 | -0.0029155 | -0.0072582 | 0.031414 | 5796.9584 | 259.2728 | 3372709.9536 | 0.52798 | 0.041571 | -0.0029145 | -0.0072605 | 0.031395 | 5796.9594 | 259.2728 | 6.2561
| 25 | 2023-11-09 17:02 | 181.7847 | 0.96529 | 19.9575 | 36.1371 | 6.2564 | 87.6078 | 1.5504e-05 | 113760 | 625.7952 | 100.056 | 6.258 | 0.024621 | 19.9575 | 37.1024 | 87.6097 | 231.7843 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 6.2645 | 0.12995 | 6.2564 | [6 empty] | 6.2744 | 0.2877 | 6.2564 | [6 empty] | 19.9526 | 35.9145 | 259.1579 | 86.8815 | 1.5493e-05 | 75885 | 417.4443 | 3372661.7655 | 5798.78 | 259.1579 | 0.52797 | 0.041571 | -0.0029154 | -0.0072611 | 0.031395 | 5796.9595 | 259.2393 | 3373720.9367 | 0.52811 | 0.041554 | -0.0029144 | -0.0072634 | 0.031377 | 5796.9605 | 259.2392 | 6.2564
| 26 | 2023-11-09 17:06 | 181.767 | 0.96524 | 19.9391 | 36.1294 | 6.2564 | 87.5956 | 1.5503e-05 | 113723 | 625.6525 | 100.0336 | 6.2565 | 0.0022074 | 19.9391 | 37.0947 | 87.5977 | 231.7758 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 6.2631 | 0.10745 | 6.2564 | [6 empty] | 6.273 | 0.26517 | 6.2564 | [6 empty] | 19.9308 | 35.9067 | 259.1865 | 86.8715 | 1.5492e-05 | 75886 | 417.4904 | 3372921.6969 | 5798.78 | 259.1865 | 0.52801 | 0.041567 | -0.0029154 | -0.0072626 | 0.031389 | 5796.9598 | 259.2679 | 3373980.7597 | 0.52814 | 0.04155 | -0.0029144 | -0.0072649 | 0.031371 | 5796.9609 | 259.2678 | 6.2564
| 27 | 2023-11-09 17:12 | 183.1039 | 0.96523 | 20.1186 | 36.1172 | 4.5985 | 87.4268 | 1.5509e-05 | 84186 | 459.7718 | 100.0313 | 4.5977 | -0.016499 | 20.1186 | 37.0824 | 87.4284 | 231.9644 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 4.6026 | 0.088881 | 4.5985 | 4.5962 | 0.11461 | 0.032922 | 0.10017 | 0.25 | 0.26932 | 4.6098 | 0.24632 | 4.5985 | 4.5962 | 0.27213 | 0.033014 | 0.10045 | 0.25 | 0.26943 | 19.9886 | 36.0026 | 189.8768 | 87.1439 | 1.5499e-05 | 56002 | 305.8483 | 2477614.3592 | 5798.78 | 189.8768 | 0.39403 | 0.05727 | -0.0040969 | -0.0053233 | 0.04785 | 5796.0053 | 189.9677 | 2478800.467 | 0.39424 | 0.057248 | -0.0040947 | -0.0053259 | 0.047827 | 5796.0066 | 189.9677 | 4.5985
| 28 | 2023-11-09 17:15 | 181.7968 | 0.96525 | 20.0911 | 36.1097 | 4.5961 | 87.4215 | 1.5508e-05 | 83554 | 459.601 | 100.0458 | 4.596 | -0.0020695 | 20.0911 | 37.075 | 87.4236 | 231.9469 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 4.6009 | 0.10324 | 4.5961 | [6 empty] | 4.6081 | 0.26073 | 4.5961 | [6 empty] | 19.9998 | 35.9952 | 189.8486 | 87.1118 | 1.5499e-05 | 55594 | 305.8029 | 2476317.9999 | 5798.78 | 189.8486 | 0.39381 | 0.057295 | -0.0040991 | -0.0053216 | 0.047874 | 5796.0039 | 189.9396 | 2477504.0847 | 0.39401 | 0.057272 | -0.0040969 | -0.0053242 | 0.047851 | 5796.0052 | 189.9395 | 4.5961
| 29 | 2023-11-09 17:19 | 181.7746 | 0.96522 | 20.0297 | 36.0938 | 4.594 | 87.4133 | 1.5505e-05 | 83546 | 459.6132 | 100.0944 | 4.5961 | 0.046525 | 20.0297 | 37.059 | 87.4155 | 231.9067 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 4.601 | 0.15171 | 4.594 | [6 empty] | 4.6082 | 0.30933 | 4.594 | [6 empty] | 20.0032 | 35.9795 | 189.8787 | 87.058 | 1.5498e-05 | 55596 | 305.8513 | 2475275.9696 | 5798.78 | 189.8787 | 0.39362 | 0.057315 | -0.0041009 | -0.0053215 | 0.047892 | 5796.0028 | 189.9697 | 2476462.0055 | 0.39383 | 0.057292 | -0.0040987 | -0.005324 | 0.047869 | 5796.0042 | 189.9696 | 4.594
| 30 | 2023-11-09 17:25 | 181.7776 | 0.96518 | 20.1251 | 36.0813 | 3.1503 | 87.3043 | 1.5508e-05 | 57237 | 314.8737 | 100.0135 | 3.1487 | -0.048841 | 20.1251 | 37.0465 | 87.3059 | 232.0142 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 3.152 | 0.056295 | 3.1503 | 3.1498 | 0.040177 | 0.018458 | 0.056163 | 0.25 | 0.25623 | 3.157 | 0.21362 | 3.1503 | 3.1498 | 0.19744 | 0.018511 | 0.056322 | 0.25 | 0.25627 | 20.0174 | 36.0353 | 129.9306 | 87.2306 | 1.5501e-05 | 38044 | 209.2887 | 1696795.1319 | 5798.78 | 129.9306 | 0.22963 | 0.072309 | -0.0063079 | -0.0036431 | 0.062358 | 5795.164 | 130.0117 | 1697853.873 | 0.2299 | 0.072289 | -0.0063034 | -0.0036454 | 0.06234 | 5795.165 | 130.0117 | 3.1503
| 31 | 2023-11-09 17:29 | 182.978 | 0.9652 | 20.1552 | 36.0803 | 3.1499 | 87.2795 | 1.5509e-05 | 57601 | 314.7974 | 100.0014 | 3.148 | -0.060961 | 20.1552 | 37.0455 | 87.2812 | 232.0443 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 3.1513 | 0.044195 | 3.1499 | [6 empty] | 3.1562 | 0.20146 | 3.1499 | [6 empty] | 20.0544 | 36.0347 | 129.957 | 87.2023 | 1.5503e-05 | 38303 | 209.3312 | 1696427.7194 | 5798.78 | 129.957 | 0.22954 | 0.072315 | -0.0063085 | -0.0036427 | 0.062364 | 5795.1636 | 130.0381 | 1697486.344 | 0.22981 | 0.072296 | -0.0063039 | -0.003645 | 0.062347 | 5795.1646 | 130.0381 | 3.1499
| 32 | 2023-11-09 17:33 | 181.7953 | 0.96522 | 20.163 | 36.0806 | 3.1493 | 87.2749 | 1.551e-05 | 57204 | 314.6616 | 99.9772 | 3.1466 | -0.085103 | 20.163 | 37.0458 | 87.2766 | 232.0515 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 3.1499 | 0.02004 | 3.1493 | [6 empty] | 3.1549 | 0.17725 | 3.1493 | [6 empty] | 20.0338 | 36.0346 | 129.9112 | 87.2165 | 1.5502e-05 | 38042 | 209.2573 | 1696198.2917 | 5798.78 | 129.9112 | 0.22948 | 0.07232 | -0.00631 | -0.003642 | 0.062368 | 5795.1634 | 129.9923 | 1697256.8302 | 0.22975 | 0.0723 | -0.0063055 | -0.0036443 | 0.06235 | 5795.1644 | 129.9922 | 3.1493
| 33 | 2023-11-09 17:40 | 183.2497 | 0.96517 | 19.9712 | 36.0261 | 1.5469 | 87.2336 | 1.5499e-05 | 28388 | 154.9143 | 100.2127 | 1.5491 | 0.14699 | 19.9712 | 36.9913 | 87.2352 | 231.9323 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 1.5508 | 0.2518 | 1.5469 | 1.5468 | 0.2336 | 0.015811 | 0.048109 | 0.25 | 0.25459 | 1.5532 | 0.40954 | 1.5469 | 1.5468 | 0.39129 | 0.015863 | 0.048265 | 0.25 | 0.25462 | 19.9174 | 36.0267 | 63.7659 | 87.2736 | 1.5497e-05 | 18822 | 102.7123 | 833385.8234 | 5798.78 | 63.7659 | -0.079154 | 0.082345 | -0.015034 | -0.0017892 | 0.065522 | 5794.9805 | 63.8077 | 833932.2321 | -0.078869 | 0.082347 | -0.015021 | -0.0017904 | 0.065535 | 5794.9798 | 63.8077 | 1.5469
| 34 | 2023-11-09 17:44 | 181.7918 | 0.96522 | 20.0389 | 36.0351 | 1.5475 | 87.2144 | 1.5502e-05 | 28165 | 154.93 | 100.1839 | 1.5493 | 0.11827 | 20.0389 | 37.0004 | 87.2166 | 231.9867 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 1.5509 | 0.22319 | 1.5475 | [6 empty] | 1.5534 | 0.38081 | 1.5475 | [6 empty] | 19.9031 | 36.0356 | 63.7616 | 87.3134 | 1.5497e-05 | 18671 | 102.7054 | 833717.7566 | 5798.78 | 63.7616 | -0.078981 | 0.082346 | -0.015029 | -0.0017895 | 0.065528 | 5794.9802 | 63.8034 | 834264.4334 | -0.078696 | 0.082347 | -0.015016 | -0.0017907 | 0.065541 | 5794.9794 | 63.8034 | 1.5475
| 35 | 2023-11-09 17:47 | 181.7634 | 0.96532 | 19.9633 | 36.0146 | 1.5462 | 87.2017 | 1.5498e-05 | 28138 | 154.8057 | 100.1868 | 1.5481 | 0.12113 | 19.9633 | 36.9799 | 87.2036 | 231.9385 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 1.5497 | 0.22583 | 1.5462 | [6 empty] | 1.5521 | 0.38352 | 1.5462 | [6 empty] | 19.8929 | 36.0156 | 63.7511 | 87.2552 | 1.5496e-05 | 18665 | 102.6885 | 833097.815 | 5798.78 | 63.7511 | -0.079304 | 0.082345 | -0.015042 | -0.001789 | 0.065514 | 5794.981 | 63.7929 | 833643.9676 | -0.079019 | 0.082346 | -0.015029 | -0.0017901 | 0.065527 | 5794.9803 | 63.7929 | 1.5462
| 36 | 2023-11-09 18:00 | 61.765 | 0.96545 | 19.8999 | 36.0203 | #NUM! | 87.2666 | 1.5496e-05 | 0 | 0 | NaN | 0 | #NUM! | 19.8999 | 36.9858 | 87.2683 | 231.8701 | 1.2686 | 92.65 | 44.7 | 0.02235 | 1 | 0.99843 | 1.0026 | 0 | #NUM! | #NUM! | [6 empty] | 0 | #NUM! | #NUM! | [6 empty] | 19.8228 | 36.0364 | 0 | 87.3746 | 1.5494e-05 | 0 | 0 | 0 | 5798.78 | 0 | #NUM! | #NUM! | #DIV/0! | 0 | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM! | #NUM!
(End of Results Table)
0.018267
Target Type | Testing.TestPoints.TestPoint | Testing.TestPoints.TestPoint | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc | Testing.Measurements.MeasurementCalcs.MeasurementCalc
Target Attribute | Number | Timestamp | 10 | 7 | 24 | 21 | 34 | 48 | 94 | 26 | 18 | 19 | 16 | 31
Target Setup Z Order | -1 | -1 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89 | 89
Display Units | - | - | s | bar | °C | barg | kg/s | kg/m³ | Pa s | pulse | Hz | pulse/kg | kg/s | %
Format |  |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0