## Data
[82 empty] | (-) | (%)
NEL | [19 empty] | FORCE | [20 empty] | DNV | [39 empty] | Total test points | 35
[81 empty] | Above 1 | 11 | 31.4286
[81 empty] | Above 1.2 | 9 | 25.7143
[81 empty] | Between 1 and 1.2 | 2 | 5.7143
[81 empty] | Below equal 1 | 24 | 68.5714
Test Number | T03973 | [15 empty] | 0.35905 | [8 empty] | Calculated by NEL with REFPROP v10 for pure CO2 | [11 empty] | 0.15119 | [19 empty] | 0.23396
Turbine meter | [19 empty] | Turbine meter | [20 empty] | Turbine meter | [19 empty] | NEL |  | FORCE |  | DNV | [6 empty] | NEL |  |  | FORCE |  |  | DNV
Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | viscosity | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | Reynolds average | Q ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
- | - | - | m3/h | bar a | °C | kg/m³ | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | - | - | - | m3/h | bar a | °C | kg/m³ | Pa s | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | - | - | - | m3/h | bar a | °C | kg/m³ | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | [6 empty] | - | m3/h | % |  |  |  |  | % |  |  | % |  |  | %
CO2 | 2024-07-10 10:02 | 4 | 398.0389 | 31.4752 | 20.0381 | 70.3042 | 6383750.4365 | 399.676 | 0.41128 | 0.35 | 0.43579 | 0.021301 | 3 | 4.3027 | 0.052914 | 0.35 | 0.35398 | 397.9555 | 6382062.5989 | CO2 | 2025-03-11 |  | 323.2249 | 31.0354 | 21.8071 | 68.141 | 1.5339e-05 | 4999634.9009 | 323.296 | 0.022002 | 0.15 | 0.029002 | 0.0063611 | 3 | 4.3027 | 0.015802 | 0.15 | 0.15083 | 323.2779 | 5013679.179 | CO2 | 2025-01-13 11:51 | 1.1 | 386.86 | 29.3686 | 24.6561 | 62.493 | 5759467.7451 | 387.2565 | 0.10249 | 0.21791 | 0.1039 | 0.0085515 | 3 | 4.3027 | 0.021243 | 0.21796 | 0.21899 | 386.7582 | 5761967.6687 | 3.478 | 7.9809 | 1.2748 | 43.9566 | 2.1665 | 20.852 | 5719236.4822 | 369.3305 | 0.095059 | 72.7895 | 0.11721 | 0.34073 | 0.33401 | 1.0201 | -0.066057 | 0.094929 | 0.69586 | 0.0088392 | 0.18498 | 0.047784
CO2 | 2024-07-10 10:12 | 5 | 397.883 | 31.4864 | 20.0916 | 70.3106 | 6380661.4545 | 399.6727 | 0.44981 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 323.3258 | 31.059 | 21.6941 | 68.474 | 1.5336e-05 | 5026947.4422 | 323.4371 | 0.034428 | 0.15 | [9 empty] | CO2 | 2025-01-13 11:53 | 1.2 | 386.7977 | 29.369 | 24.5673 | 62.5228 | 5762155.9684 | 387.1696 | 0.09614 | 0.21786
CO2 | 2024-07-10 10:22 | 6 | 397.9446 | 31.4898 | 20.103 | 70.3149 | 6381775.9056 | 399.7206 | 0.44629 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 323.283 | 31.0879 | 21.9254 | 68.3656 | 1.5348e-05 | 5014455.1939 | 323.3819 | 0.030575 | 0.15 | [9 empty] | CO2 | 2025-01-13 11:59 | 1.3 | 386.6169 | 29.3662 | 24.4627 | 62.5601 | 5764279.2928 | 387.054 | 0.11307 | 0.2181
CO2 | 2024-07-10 10:48 | 7 | 278.7836 | 31.4841 | 20.2379 | 70.2336 | 4463932.9327 | 279.8617 | 0.38674 | 0.35 | 0.40698 | 0.03878 | 3 | 4.3027 | 0.096335 | 0.35 | 0.36302 | 278.7041 | 4463186.6972 | CO2 | 2025-03-11 |  | 282.226 | 31.1447 | 22.0052 | 68.3296 | 1.5351e-05 | 4379579.1631 | 282.6567 | 0.1526 | 0.15 | 0.1541 | 0.0072008 | 3 | 4.3027 | 0.017888 | 0.15 | 0.15106 | 282.1825 | 4385119.6786 | CO2 | 2025-01-13 12:18 | 3.1 | 272.8754 | 31.0155 | 25.9289 | 66.3141 | 4258595.5765 | 273.1004 | 0.082428 | 0.21787 | 0.083696 | 0.0051381 | 3 | 4.3027 | 0.012764 | 0.21702 | 0.21739 | 272.8745 | 4258214.5405 | 3.0883 | 7.5884 | 6.7528 | 43.8213 | 1.771 | 21.1598 | 4368840.3054 | 277.9204 | 0.16001 | 72.5694 | 0.11739 | 0.24696 | 0.34351 | 0.71894 | -0.005915 | 0.095079 | 0.062211 | -0.076317 | 0.18297 | 0.41709
CO2 | 2024-07-10 10:59 | 8 | 278.7704 | 31.4641 | 20.171 | 70.2085 | 4463213.1013 | 279.8367 | 0.3825 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 282.1975 | 31.1606 | 22.1586 | 68.5405 | 1.5362e-05 | 4389419.0073 | 282.6145 | 0.14777 | 0.15 | [9 empty] | CO2 | 2025-01-13 12:21 | 3.2 | 272.8699 | 31.0012 | 25.8781 | 66.294 | 4257935.0584 | 273.1137 | 0.08935 | 0.21672
CO2 | 2024-07-10 11:08 | 9 | 278.5583 | 31.4262 | 19.9623 | 70.2 | 4462414.0575 | 279.8165 | 0.45169 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 282.1241 | 31.173 | 22.2214 | 68.5134 | 1.5365e-05 | 4386360.8654 | 282.5809 | 0.16193 | 0.15 | [9 empty] | CO2 | 2025-01-13 12:23 | 3.3 | 272.8782 | 30.997 | 25.8526 | 66.2905 | 4258112.9866 | 273.0947 | 0.079311 | 0.21647
CO2 | 2024-07-10 11:20 | 10 | 159.8768 | 31.4195 | 20.0093 | 70.1581 | 2559330.684 | 160.3776 | 0.31327 | 0.35 | 0.32784 | 0.019271 | 3 | 4.3027 | 0.047872 | 0.35 | 0.35326 | 159.8314 | 2557302.8665 | CO2 | 2025-03-11 |  | 160.4544 | 31.174 | 21.7005 | 68.749 | 1.534e-05 | 2511503.9764 | 160.9873 | 0.3321 | 0.15 | 0.33824 | 0.0070152 | 3 | 4.3027 | 0.017427 | 0.15 | 0.15101 | 160.4511 | 2511930.8241 | CO2 | 2025-01-13 13:09 | 4.1 | 158.6008 | 31.7114 | 26.2847 | 68.0087 | 2526044.137 | 158.7618 | 0.10149 | 0.22897 | 0.11521 | 0.013796 | 3 | 4.3027 | 0.03427 | 0.22465 | 0.22725 | 158.5711 | 2524249.1274 | 2.6271 | 8.0134 | 14.8326 | 43.8526 | 2.231 | 19.3642 | 2531160.9393 | 159.6179 | 0.27644 | 71.2301 | 0.11849 | 0.051399 | 0.3328 | 0.15445 | 0.061801 | 0.09362 | 0.66013 | -0.16123 | 0.19391 | 0.83143
CO2 | 2024-07-10 11:31 | 11 | 159.7903 | 31.4133 | 20.0489 | 70.1216 | 2556353.2951 | 160.349 | 0.34969 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 160.4481 | 31.1716 | 21.6852 | 68.7536 | 1.534e-05 | 2511993.2114 | 161.003 | 0.34588 | 0.15 | [9 empty] | CO2 | 2025-01-13 13:11 | 4.2 | 158.5507 | 31.6678 | 26.1831 | 67.9347 | 2523617.9697 | 158.7554 | 0.12908 | 0.22383
CO2 | 2024-07-10 11:41 | 12 | 159.8272 | 31.4035 | 20.039 | 70.0982 | 2556224.6203 | 160.3395 | 0.32055 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 160.4509 | 31.179 | 21.6748 | 68.7627 | 1.5339e-05 | 2512295.2845 | 160.9912 | 0.33673 | 0.15 | [9 empty] | CO2 | 2025-01-13 13:13 | 4.3 | 158.5619 | 31.6376 | 26.0917 | 67.8922 | 2523085.2754 | 158.7444 | 0.11506 | 0.22115
CO2 | 2024-07-10 12:06 | 14 | 80.4381 | 31.3871 | 20.0195 | 70.0606 | 1285934.492 | 80.6345 | 0.2441 | 0.35 | 0.24072 | 0.0078865 | 3 | 4.3027 | 0.019591 | 0.35 | 0.35055 | 80.438 | 1285436.4355 | CO2 | 2025-03-11 |  | 80.0733 | 31.1941 | 21.5709 | 68.8083 | 1.5334e-05 | 1254571.0977 | 80.3167 | 0.30398 | 0.15 | 0.30164 | 0.0022779 | 3 | 4.3027 | 0.0056585 | 0.15 | 0.15011 | 80.0735 | 1255520.1157 | CO2 | 2025-01-13 13:31 | 5.1 | 83.3558 | 31.4329 | 25.1309 | 67.7693 | 1327886.1033 | 83.3212 | -0.041494 | 0.29059 | 0.0014191 | 0.045054 | 3 | 4.3027 | 0.11192 | 0.26678 | 0.2893 | 83.2964 | 1326328.3567 | 1.9589 | 8.1378 | 13.3873 | 44.3813 | 0.016956 | 11.9481 | 1289094.9693 | 81.2693 | 0.23831 | 64.4672 | 0.12455 | 0.0024127 | 0.32768 | 0.007363 | 0.063332 | 0.083787 | 0.75587 | -0.23689 | 0.26112 | 0.90721
CO2 | 2024-07-10 12:25 | 15 | 80.4357 | 31.3837 | 20.0488 | 70.0368 | 1285359.5582 | 80.6339 | 0.24636 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 80.0737 | 31.1779 | 21.4682 | 68.8541 | 1.533e-05 | 1255738.0507 | 80.3135 | 0.29942 | 0.15 | [9 empty] | CO2 | 2025-01-13 13:33 | 5.2 | 83.297 | 31.34 | 24.8358 | 67.6454 | 1325961.1949 | 83.2949 | -0.0025943 | 0.26801
CO2 | 2024-07-10 12:45 | 16 | 80.4401 | 31.3805 | 20.069 | 70.0178 | 1285015.2562 | 80.6265 | 0.23171 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 80.0736 | 31.1711 | 21.4307 | 68.8724 | 1.5328e-05 | 1256251.1986 | 80.3151 | 0.30153 | 0.15 | [9 empty] | CO2 | 2025-01-13 13:35 | 5.3 | 83.2365 | 31.2742 | 24.5439 | 67.5913 | 1325137.772 | 83.2767 | 0.048346 | 0.24173
CO2 | 2024-07-10 13:06 | 17 | 40.3212 | 31.3729 | 20.0135 | 70.0225 | 644282.5445 | 40.3757 | 0.13511 | 0.35 | 0.12935 | 0.0098121 | 3 | 4.3027 | 0.024375 | 0.35 | 0.35085 | 40.3215 | 644310.696 | CO2 | 2025-03-11 |  | 40.018 | 31.0521 | 20.6953 | 68.9266 | 1.529e-05 | 629016.9339 | 40.0838 | 0.16456 | 0.15 | 0.1508 | 0.012429 | 3 | 4.3027 | 0.030876 | 0.15 | 0.15314 | 40.0194 | 628850.3924 | CO2 | 2025-01-13 14:11 | 8.1 | 34.1293 | 31.5052 | 24.5356 | 68.2029 | 547460.0911 | 33.9799 | -0.43754 | 0.22635 | -0.4406 | 0.0041312 | 3 | 4.3027 | 0.010262 | 0.22625 | 0.22648 | 34.1365 | 547381.4104 | 1.0508 | 8.1239 | 6.4296 | 42.6379 | -8.5897 | 19.4954 | 606847.4996 | 38.1591 | -0.015789 | 70.2571 | 0.1193 | 0.14513 | 0.32994 | 0.43988 | 0.16659 | 0.09602 | 1.7349 | -0.42482 | 0.19251 | 2.2067
CO2 | 2024-07-10 13:26 | 18 | 40.326 | 31.3615 | 19.9591 | 70.0161 | 644421.0062 | 40.3736 | 0.11802 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 40.0235 | 31.0548 | 20.7179 | 68.9174 | 1.5291e-05 | 628829.8995 | 40.0797 | 0.14038 | 0.15 | [9 empty] | CO2 | 2025-01-13 14:14 | 8.2 | 34.1381 | 31.5106 | 24.5978 | 68.1932 | 547449.4494 | 33.9861 | -0.4453 | 0.22599
CO2 | 2024-07-10 13:46 | 19 | 40.3172 | 31.3545 | 19.94 | 70.0053 | 644228.5373 | 40.3716 | 0.13491 | 0.35 | [9 empty] | CO2 | 2025-03-11 |  | 40.0166 | 31.0585 | 20.7253 | 68.9128 | 1.5291e-05 | 628704.3438 | 40.0756 | 0.14745 | 0.15 | [9 empty] | CO2 | 2025-01-13 14:16 | 8.3 | 34.1421 | 31.507 | 24.6411 | 68.1625 | 547234.6908 | 33.9922 | -0.43897 | 0.22641
CO2 | 2024-07-10 14:08 | 20 | 20.4295 | 31.3539 | 19.8701 | 70.0368 | 326657.4838 | 20.4678 | 0.18768 | 0.35 | 0.28659 | 0.086045 | 3 | 4.3027 | 0.21375 | 0.35 | 0.41011 | 20.3372 | 325275.471 | [21 empty] | CO2 | 2025-01-13 14:20 | 9.1 | 24.4537 | 31.5271 | 24.6541 | 68.206 | 392132.9469 | 24.341 | -0.46101 | 0.22302 | -0.50818 | 0.04597 | 3 | 4.3027 | 0.1142 | 0.22324 | 0.25075 | 24.4521 | 392152.0141 | 1.704 | 5.9457 |  |  | -8.0823 | 15.9042 | 358713.7425 | 22.3947 | -0.29191 | 21.85 | 0.21393 | 0.5785 | 0.34989 | 1.6534 | 0.29191 |  |  | -0.21627 | 0.1308 | 1.6534
CO2 | 2024-07-10 14:27 | 21 | 20.3153 | 31.3652 | 19.868 | 70.0702 | 324979.2832 | 20.3819 | 0.32796 | 0.35 | [30 empty] | CO2 | 2025-01-13 14:23 | 9.2 | 24.452 | 31.5318 | 24.6383 | 68.2219 | 392194.7424 | 24.3272 | -0.51069 | 0.2232
CO2 | 2024-07-10 14:46 | 22 | 20.2669 | 31.3759 | 19.9154 | 70.0783 | 324189.6459 | 20.3367 | 0.34414 | 0.35 | [30 empty] | CO2 | 2025-01-13 14:25 | 9.3 | 24.4506 | 31.5255 | 24.6213 | 68.2099 | 392128.353 | 24.3154 | -0.55285 | 0.2235
Test Number | T03981 | [24 empty] | Calculated by NEL with REFPROP v10 for pure CO2
Turbine meter | [19 empty] | Turbine meter | [20 empty] | Turbine meter | [19 empty] | NEL |  | FORCE |  | DNV | [6 empty] | NEL |  |  | FORCE |  |  | DNV
Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | viscosity | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | xi/u2(xi) | 1/u2(xi) | Reynolds average | Q ref average | KCRV y | 1/u2(y) | u(y) | di | u(di) | En (%) | di | u(di) | En (%) | di | u(di) | En (%)
- | - | - | m3/h | bar a | °C | kg/m³ | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | - | - | - | m3/h | bar a | °C | kg/m³ | Pa s | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | - | - | - | m3/h | bar a | °C | kg/m³ | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | - | [6 empty] | - | m3/h | % |  |  |  |  | % |  |  | % |  |  | %
CO2 | 2024-07-11 09:49 | 2 | 398.1693 | 21.3807 | 20.0634 | 44.0073 | 4075439.302 | 399.9076 | 0.43657 | 0.35 | 0.44502 | 0.0094857 | 3 | 4.3027 | 0.023564 | 0.35 | 0.35079 | 398.1154 | 4076958.4786 | CO2 | 2025-03-07 |  | 342.5586 | 20.7981 | 22.106 | 42.1854 | 1.5067e-05 | 3344876.1449 | 343.1909 | 0.18459 | 0.15 | 0.18178 | 0.0079215 | 3 | 4.3027 | 0.019678 | 0.15 | 0.15129 | 342.5664 | 3342471.4535 | CO2 | 2025-01-14 09:03 | 11.1 | 386.6702 | 21.1023 | 20.8204 | 43.1057 | 4166484.5903 | 386.7205 | 0.012995 | 0.26191 | -0.0039622 | 0.016077 | 3 | 4.3027 | 0.039937 | 0.237 | 0.24035 | 386.6044 | 4172588.2835 | 3.6164 | 8.1264 | 7.9425 | 43.6925 | -0.06859 | 17.3112 | 3864006.0719 | 375.7621 | 0.16621 | 69.1302 | 0.12027 | 0.27881 | 0.32953 | 0.84608 | 0.015568 | 0.09177 | 0.16965 | -0.17018 | 0.20809 | 0.81781
CO2 | 2024-07-11 09:59 | 3 | 398.1174 | 21.3844 | 20.0074 | 44.0289 | 4077627.3384 | 399.8819 | 0.44321 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 342.5118 | 20.7982 | 22.2174 | 42.1618 | 1.5073e-05 | 3341451.3415 | 343.1554 | 0.18791 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:06 | 11.2 | 386.6286 | 21.1403 | 20.8374 | 43.1914 | 4173560.9979 | 386.6058 | -0.0058989 | 0.22704
CO2 | 2024-07-11 10:09 | 4 | 398.0595 | 21.3915 | 20.0294 | 44.0409 | 4077808.7953 | 399.8718 | 0.45528 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 342.6289 | 20.8094 | 22.3261 | 42.1647 | 1.5078e-05 | 3341086.8742 | 343.2211 | 0.17284 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:08 | 11.3 | 386.5144 | 21.1571 | 20.7717 | 43.2438 | 4177719.2623 | 386.4411 | -0.018983 | 0.22206
CO2 | 2024-07-11 10:20 | 5 | 278.939 | 21.3944 | 20.0943 | 44.0329 | 2856376.2414 | 280.06 | 0.40186 | 0.35 | 0.41398 | 0.016098 | 3 | 4.3027 | 0.039991 | 0.35 | 0.35228 | 278.8697 | 2855489.0265 | CO2 | 2025-03-07 |  | 281.6757 | 20.858 | 22.4179 | 42.2577 | 1.5084e-05 | 2751951.4162 | 282.183 | 0.18012 | 0.15 | 0.17861 | 0.0032175 | 3 | 4.3027 | 0.0079926 | 0.15 | 0.15021 | 281.6449 | 2751095.9861 | CO2 | 2025-01-14 09:24 | 13.1 | 277.2986 | 22.3159 | 20.803 | 45.984 | 3172885.8436 | 277.4439 | 0.052395 | 0.22011 | 0.062426 | 0.008928 | 3 | 4.3027 | 0.022178 | 0.22084 | 0.22195 | 277.2576 | 3174560.6916 | 3.3359 | 8.0581 | 7.9157 | 44.3186 | 1.2672 | 20.2996 | 2927048.568 | 279.2574 | 0.17225 | 72.6763 | 0.1173 | 0.24173 | 0.33217 | 0.72771 | 0.0063546 | 0.093831 | 0.067724 | -0.10983 | 0.18842 | 0.58289
CO2 | 2024-07-11 10:30 | 6 | 278.8598 | 21.3816 | 20.013 | 44.0212 | 2855616.3239 | 279.9971 | 0.40783 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 281.6037 | 20.8366 | 22.3334 | 42.2262 | 1.5079e-05 | 2750056.5761 | 282.1128 | 0.18079 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:27 | 13.2 | 277.241 | 22.3108 | 20.6718 | 46.0046 | 3174715.6276 | 277.4337 | 0.069504 | 0.2214
CO2 | 2024-07-11 10:50 | 8 | 278.8101 | 21.3727 | 19.9841 | 44.0067 | 2854474.5141 | 280.0153 | 0.43225 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 281.6553 | 20.8458 | 22.3538 | 42.2431 | 1.508e-05 | 2751279.966 | 282.1479 | 0.17491 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:29 | 13.3 | 277.2333 | 22.3008 | 20.5455 | 46.0105 | 3176080.6035 | 277.4145 | 0.065378 | 0.221
CO2 | 2024-07-11 11:02 | 9 | 160.0279 | 21.3765 | 20.0676 | 43.9966 | 1637543.8766 | 160.4672 | 0.27454 | 0.35 | 0.28714 | 0.017935 | 3 | 4.3027 | 0.044554 | 0.35 | 0.35282 | 159.9921 | 1637088.3353 | CO2 | 2025-03-07 |  | 160.3387 | 20.86 | 21.9563 | 42.3613 | 1.5061e-05 | 1573992.9409 | 160.7606 | 0.26313 | 0.15 | 0.25673 | 0.0098706 | 3 | 4.3027 | 0.02452 | 0.15 | 0.15199 | 160.3391 | 1573955.8968 | CO2 | 2025-01-14 09:37 | 14.1 | 161.7931 | 23.0601 | 21.1976 | 47.6742 | 1911644.4729 | 161.997 | 0.126 | 0.22224 | 0.13091 | 0.0062244 | 3 | 4.3027 | 0.015462 | 0.22058 | 0.22112 | 161.7569 | 1911519.2588 | 2.3066 | 8.0331 | 11.1132 | 43.2878 | 2.6774 | 20.452 | 1707521.1636 | 160.696 | 0.22428 | 71.7728 | 0.11804 | 0.062855 | 0.33249 | 0.18904 | 0.032449 | 0.095752 | 0.33889 | -0.093368 | 0.18698 | 0.49934
CO2 | 2024-07-11 11:12 | 10 | 160.0032 | 21.3625 | 19.9692 | 43.9857 | 1637447.8984 | 160.45 | 0.2792 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 160.3312 | 20.851 | 21.8972 | 42.3531 | 1.5058e-05 | 1573916.7822 | 160.7508 | 0.2617 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:39 | 14.2 | 161.7412 | 23.0645 | 21.1873 | 47.6874 | 1911565.6074 | 161.9642 | 0.13791 | 0.21959
CO2 | 2024-07-11 11:23 | 11 | 159.9452 | 21.3583 | 19.9848 | 43.9721 | 1636273.2307 | 160.4373 | 0.30767 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 160.3475 | 20.8456 | 21.852 | 42.3503 | 1.5056e-05 | 1573957.9673 | 160.7409 | 0.24536 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:41 | 14.3 | 161.7365 | 23.0608 | 21.1757 | 47.6813 | 1911347.696 | 161.9448 | 0.12883 | 0.21992
CO2 | 2024-07-11 11:52 | 12 | 80.0068 | 21.3531 | 19.994 | 43.9577 | 818201.5712 | 80.0065 | -0.00037182 | 0.35 | -0.029359 | 0.027614 | 3 | 4.3027 | 0.068596 | 0.35 | 0.35666 | 80.0194 | 818088.8864 | CO2 | 2025-03-07 |  | 80.0564 | 20.9406 | 20.7465 | 42.8133 | 1.5004e-05 | 797406.3276 | 80.2761 | 0.27442 | 0.15 | 0.26599 | 0.0077871 | 3 | 4.3027 | 0.019344 | 0.15 | 0.15124 | 80.0556 | 798146.8502 | CO2 | 2025-01-14 09:56 | 15.1 | 85.3928 | 23.305 | 21.2974 | 48.2464 | 1019835.7912 | 85.3359 | -0.066681 | 0.22269 | -0.084624 | 0.016499 | 3 | 4.3027 | 0.040986 | 0.22241 | 0.22615 | 85.3925 | 1019901.3994 | -0.2308 | 7.8613 | 11.6284 | 43.7174 | -1.6546 | 19.5519 | 878712.3787 | 81.8225 | 0.13697 | 71.1306 | 0.11857 | -0.16633 | 0.33637 | 0.49449 | 0.12902 | 0.093891 | 1.3741 | -0.2216 | 0.19258 | 1.1507
CO2 | 2024-07-11 12:11 | 13 | 80.0231 | 21.349 | 20.0053 | 43.9454 | 818113.7323 | 79.9972 | -0.032349 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 80.0547 | 20.9386 | 20.575 | 42.8464 | 1.4995e-05 | 798376.7727 | 80.2664 | 0.26448 | 0.15 | [9 empty] | CO2 | 2025-01-14 09:58 | 15.2 | 85.3937 | 23.3159 | 21.3352 | 48.2604 | 1019993.0775 | 85.3185 | -0.088048 | 0.22147
CO2 | 2024-07-11 12:31 | 14 | 80.0284 | 21.3455 | 20.0146 | 43.9348 | 817951.3556 | 79.9841 | -0.055356 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 80.0557 | 20.931 | 20.4806 | 42.8495 | 1.499e-05 | 798657.4502 | 80.2631 | 0.25907 | 0.15 | [9 empty] | CO2 | 2025-01-14 10:00 | 15.3 | 85.391 | 23.3214 | 21.375 | 48.2623 | 1019875.3295 | 85.3064 | -0.099142 | 0.22307
CO2 | 2024-07-11 12:52 | 15 | 40.4292 | 21.3443 | 19.971 | 43.942 | 413343.1498 | 40.3475 | -0.2019 | 0.35 | -0.23236 | 0.037906 | 3 | 4.3027 | 0.094163 | 0.35 | 0.36245 | 40.4385 | 413400.7405 | CO2 | 2025-03-07 |  | 40.0177 | 20.9203 | 20.3371 | 42.8559 | 1.4983e-05 | 398989.4805 | 40.0721 | 0.13577 | 0.15 | 0.14389 | 0.0099108 | 3 | 4.3027 | 0.02462 | 0.15 | 0.15201 | 40.0165 | 398954.2374 | CO2 | 2025-01-14 10:20 | 17.1 | 34.9941 | 23.0555 | 21.1782 | 47.669 | 413443.4963 | 34.7647 | -0.65558 | 0.23262 | -0.65561 | 0.0010833 | 3 | 4.3027 | 0.0026911 | 0.23487 | 0.23488 | 34.9909 | 413044.0046 | -1.7688 | 7.6123 | 6.2276 | 43.2786 | -11.8834 | 18.1257 | 408466.3275 | 38.482 | -0.10758 | 69.0166 | 0.12037 | -0.12478 | 0.34187 | 0.365 | 0.25147 | 0.092827 | 2.709 | -0.54803 | 0.2017 | 2.7171
CO2 | 2024-07-11 13:11 | 16 | 40.4338 | 21.3405 | 19.9521 | 43.9373 | 413375.1237 | 40.3447 | -0.22038 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 40.0156 | 20.9152 | 20.2918 | 42.854 | 1.4981e-05 | 399034.5323 | 40.0721 | 0.14097 | 0.15 | [9 empty] | CO2 | 2025-01-14 10:22 | 17.2 | 34.9905 | 23.0308 | 21.1338 | 47.619 | 413047.9266 | 34.7607 | -0.65671 | 0.23614
CO2 | 2024-07-11 13:30 | 17 | 40.4526 | 21.3419 | 19.9848 | 43.9333 | 413483.9482 | 40.3414 | -0.27481 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 40.0162 | 20.9033 | 20.2922 | 42.8258 | 1.4981e-05 | 398838.6994 | 40.0782 | 0.15494 | 0.15 | [9 empty] | CO2 | 2025-01-14 10:24 | 17.3 | 34.9882 | 23.0043 | 21.086 | 47.5651 | 412640.5908 | 34.7592 | -0.65455 | 0.23584
CO2 | 2024-07-11 13:50 | 18 | 20.8814 | 21.3414 | 19.9413 | 43.9419 | 213510.5398 | 20.8272 | -0.25956 | 0.35 | -0.2658 | 0.014711 | 3 | 4.3027 | 0.036544 | 0.35 | 0.3519 | 20.8407 | 213135.7148 | CO2 | 2025-03-07 |  | 20.0077 | 20.9021 | 20.2838 | 42.8248 | 1.498e-05 | 198991.561 | 19.996 | -0.058701 | 0.15 | -0.061548 | 0.0034124 | 3 | 4.3027 | 0.0084769 | 0.15 | 0.15024 | 20.0072 | 198924.7808 | CO2 | 2025-01-14 10:32 | 18.1 | 25.0498 | 22.8867 | 20.8928 | 47.3255 | 294210.3768 | 24.8454 | -0.81594 | 0.23326 | -0.81791 | 0.0021842 | 3 | 4.3027 | 0.0054259 | 0.23282 | 0.23289 | 25.0477 | 293929.6533 | -2.1464 | 8.0752 | -2.7268 | 44.303 | -15.0805 | 18.4378 | 235330.0496 | 21.9652 | -0.28177 | 70.816 | 0.11883 | 0.015968 | 0.33123 | 0.048207 | 0.22022 | 0.091928 | 2.3956 | -0.53614 | 0.20029 | 2.6769
CO2 | 2024-07-11 14:09 | 19 | 20.8295 | 21.3451 | 19.9298 | 43.9533 | 213042.0381 | 20.7707 | -0.2826 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 20.0076 | 20.8969 | 20.3052 | 42.8079 | 1.4981e-05 | 198886.3249 | 19.9945 | -0.06533 | 0.15 | [9 empty] | CO2 | 2025-01-14 10:34 | 18.2 | 25.0478 | 22.8587 | 20.8415 | 47.2715 | 293922.5219 | 24.843 | -0.81754 | 0.23424
CO2 | 2024-07-11 14:28 | 20 | 20.8111 | 21.3467 | 19.9391 | 43.955 | 212854.5666 | 20.7579 | -0.25524 | 0.35 | [9 empty] | CO2 | 2025-03-07 |  | 20.0062 | 20.8961 | 20.291 | 42.8091 | 1.498e-05 | 198896.4564 | 19.994 | -0.060613 | 0.15 | [9 empty] | CO2 | 2025-01-14 10:36 | 18.3 | 25.0456 | 22.8333 | 20.7973 | 47.2225 | 293656.0612 | 24.8401 | -0.82026 | 0.23097

## Data NG FORCE
FORCE
[6 empty] | Calculated by NEL with REFPROP v10 for 0.98 mol CH4, 0.013 mol N2, 0.007 mol CO2 | [11 empty] | 0.15019
Turbine Meter
Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | viscosity | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average
- | - | - | m3/h | bar a | °C | kg/m³ | Pa s | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | -
NG | 2025-03-11 |  | 401.0693 | 31.0587 | 20.5105 | 22.1019 | 1.1607e-05 | 2669933.6645 | 402.7717 | 0.42447 | 0.15 | 0.42825 | 0.0033128 | 3 | 4.3027 | 0.0082295 | 0.15 | 0.15023 | 401.1687 | 2666950.4798
NG | 2025-03-11 |  | 401.1859 | 31.0766 | 20.7244 | 22.0826 | 1.1613e-05 | 2666977.7641 | 402.9097 | 0.42969 | 0.15
NG | 2025-03-11 |  | 401.2509 | 31.106 | 20.9065 | 22.0651 | 1.1619e-05 | 2663940.0107 | 402.9787 | 0.43061 | 0.15
NG | 2025-03-11 |  | 320.6433 | 31.1411 | 20.938 | 22.0623 | 1.162e-05 | 2127555.995 | 321.9056 | 0.39367 | 0.15 | 0.38979 | 0.0035532 | 3 | 4.3027 | 0.0088266 | 0.15 | 0.15026 | 320.658 | 2127082.3501
NG | 2025-03-11 |  | 320.6793 | 31.1436 | 20.9726 | 22.0596 | 1.1621e-05 | 2127207.1352 | 321.9194 | 0.38669 | 0.15
NG | 2025-03-11 |  | 320.6513 | 31.1438 | 21.0051 | 22.0559 | 1.1622e-05 | 2126483.9201 | 321.8987 | 0.38902 | 0.15
NG | 2025-03-11 |  | 280.5467 | 31.1587 | 21.0646 | 22.1262 | 1.1626e-05 | 1865024.5462 | 281.5148 | 0.34505 | 0.15 | 0.34899 | 0.0036607 | 3 | 4.3027 | 0.0090936 | 0.15 | 0.15028 | 280.528 | 1864673.7063
NG | 2025-03-11 |  | 280.5198 | 31.1586 | 21.0853 | 22.1235 | 1.1627e-05 | 1864545.9168 | 281.5007 | 0.34966 | 0.15
NG | 2025-03-11 |  | 280.5175 | 31.1588 | 21.1014 | 22.1225 | 1.1627e-05 | 1864450.656 | 281.5057 | 0.35228 | 0.15
NG | 2025-03-11 |  | 160.1233 | 31.1572 | 20.8828 | 22.1428 | 1.162e-05 | 1065178.9478 | 160.5832 | 0.2872 | 0.15 | 0.28651 | 0.000652 | 3 | 4.3027 | 0.0016197 | 0.15 | 0.15001 | 160.1251 | 1064146.3767
NG | 2025-03-11 |  | 160.1293 | 31.1556 | 20.8641 | 22.1447 | 1.162e-05 | 1065352.0776 | 160.5872 | 0.2859 | 0.15
NG | 2025-03-11 |  | 160.1226 | 31.1491 | 20.8554 | 22.0697 | 1.1618e-05 | 1061908.1046 | 160.5812 | 0.28643 | 0.15
NG | 2025-03-11 |  | 80.0306 | 31.1498 | 20.7554 | 22.0789 | 1.1614e-05 | 530526.4027 | 80.1706 | 0.17497 | 0.15 | 0.17439 | 0.0007082 | 3 | 4.3027 | 0.0017593 | 0.15 | 0.15001 | 80.0311 | 531166.7838
NG | 2025-03-11 |  | 80.0309 | 31.1521 | 20.732 | 22.1567 | 1.1616e-05 | 532337.6899 | 80.1706 | 0.17459 | 0.15
NG | 2025-03-11 |  | 80.0317 | 31.1459 | 20.7265 | 22.0816 | 1.1614e-05 | 530636.2589 | 80.1706 | 0.1736 | 0.15
NG | 2025-03-11 |  | 40.0106 | 31.1291 | 20.6877 | 22.0853 | 1.1612e-05 | 264948.391 | 40.0184 | 0.019423 | 0.15 | 0.02005 | 0.0043535 | 3 | 4.3027 | 0.010815 | 0.15 | 0.15039 | 40.0109 | 265000.5623
NG | 2025-03-11 |  | 40.0116 | 31.1272 | 20.669 | 22.0872 | 1.1612e-05 | 265005.2783 | 40.0215 | 0.024683 | 0.15
NG | 2025-03-11 |  | 40.0104 | 31.1276 | 20.6318 | 22.0908 | 1.161e-05 | 265048.0177 | 40.0169 | 0.016043 | 0.15
NG | 2025-03-11 |  | 20.0042 | 31.1248 | 20.614 | 22.0927 | 1.161e-05 | 131233.0773 | 19.8108 | -0.96691 | 0.15 | -0.9717 | 0.0078032 | 3 | 4.3027 | 0.019384 | 0.15 | 0.15125 | 20.0052 | 131257.2517
NG | 2025-03-11 |  | 20.0048 | 31.1172 | 20.5773 | 22.0954 | 1.1609e-05 | 131245.8366 | 19.8086 | -0.98071 | 0.15
NG | 2025-03-11 |  | 20.0065 | 31.1171 | 20.559 | 22.0973 | 1.1608e-05 | 131292.8413 | 19.813 | -0.9675 | 0.15
[6 empty] | Calculated by NEL with REFPROP v10 for 0.98 mol CH4, 0.013 mol N2, 0.007 mol CO2 | [11 empty] | 0.15022
Turbine Meter
Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | viscosity | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average
- | - | - | m3/h | bar a | °C | kg/m³ | Pa s | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | -
NG | 2025-03-13 |  | 400.9943 | 21.019 | 20.5619 | 14.6482 | 1.1401e-05 | 1801162.3331 | 402.7027 | 0.42605 | 0.15 | 0.4317 | 0.0057517 | 3 | 4.3027 | 0.014288 | 0.15 | 0.15068 | 400.9229 | 1799641.7987
NG | 2025-03-13 |  | 400.9314 | 21.0298 | 20.6777 | 14.6414 | 1.1405e-05 | 1799512.8882 | 402.6614 | 0.4315 | 0.15
NG | 2025-03-13 |  | 400.8431 | 21.0323 | 20.7628 | 14.6368 | 1.1407e-05 | 1798250.1748 | 402.597 | 0.43754 | 0.15
NG | 2025-03-13 |  | 340.7388 | 21.0507 | 20.7303 | 14.7109 | 1.1408e-05 | 1534774.4637 | 341.9058 | 0.34247 | 0.15 | 0.34005 | 0.0021089 | 3 | 4.3027 | 0.0052388 | 0.15 | 0.15009 | 340.7597 | 1534692.8221
NG | 2025-03-13 |  | 340.7582 | 21.0512 | 20.7422 | 14.7103 | 1.1408e-05 | 1534705.8558 | 341.9137 | 0.3391 | 0.15
NG | 2025-03-13 |  | 340.7822 | 21.0506 | 20.7608 | 14.7092 | 1.1409e-05 | 1534598.1469 | 341.936 | 0.33859 | 0.15
NG | 2025-03-13 |  | 280.4549 | 21.0668 | 20.7265 | 14.7109 | 1.1408e-05 | 1263105.2336 | 281.3853 | 0.33176 | 0.15 | 0.3287 | 0.0030574 | 3 | 4.3027 | 0.0075951 | 0.15 | 0.15019 | 280.4514 | 1263079.7089
NG | 2025-03-13 |  | 280.4804 | 21.0671 | 20.7251 | 14.7109 | 1.1408e-05 | 1263181.8073 | 281.4024 | 0.3287 | 0.15
NG | 2025-03-13 |  | 280.419 | 21.0649 | 20.7235 | 14.7115 | 1.1408e-05 | 1262952.0857 | 281.3321 | 0.32565 | 0.15
NG | 2025-03-13 |  | 160.1111 | 21.0718 | 20.5937 | 14.719 | 1.1404e-05 | 721115.9732 | 160.4925 | 0.23824 | 0.15 | 0.23833 | 0.0018302 | 3 | 4.3027 | 0.0045465 | 0.15 | 0.15007 | 160.1096 | 721256.7919
NG | 2025-03-13 |  | 160.1079 | 21.0703 | 20.5518 | 14.7213 | 1.1402e-05 | 721311.6002 | 160.4925 | 0.24021 | 0.15
NG | 2025-03-13 |  | 160.1099 | 21.0735 | 20.541 | 14.7218 | 1.1402e-05 | 721342.8022 | 160.4886 | 0.23655 | 0.15
NG | 2025-03-13 |  | 80.0249 | 21.0781 | 20.4941 | 14.7247 | 1.14e-05 | 360192.4133 | 80.1106 | 0.10698 | 0.15 | 0.10902 | 0.0019769 | 3 | 4.3027 | 0.004911 | 0.15 | 0.15008 | 80.0239 | 360292.6192
NG | 2025-03-13 |  | 80.0232 | 21.0734 | 20.4519 | 14.727 | 1.1399e-05 | 360290.1705 | 80.1106 | 0.10916 | 0.15
NG | 2025-03-13 |  | 80.0234 | 21.0682 | 20.4082 | 14.7293 | 1.1398e-05 | 360395.2738 | 80.1122 | 0.11093 | 0.15
NG | 2025-03-13 |  | 40.0084 | 21.068 | 20.3616 | 14.7322 | 1.1396e-05 | 179682.1267 | 39.9279 | -0.2012 | 0.15 | -0.19928 | 0.0018022 | 3 | 4.3027 | 0.004477 | 0.15 | 0.15007 | 40.0082 | 179708.8241
NG | 2025-03-13 |  | 40.0091 | 21.0603 | 20.3473 | 14.7328 | 1.1396e-05 | 179701.2226 | 39.9295 | -0.19903 | 0.15
NG | 2025-03-13 |  | 40.007 | 21.061 | 20.3059 | 14.7351 | 1.1394e-05 | 179743.1228 | 39.9279 | -0.19762 | 0.15
NG | 2025-03-13 |  | 20.0049 | 21.0368 | 20.0045 | 14.6804 | 1.1382e-05 | 88993.8072 | 19.8217 | -0.91605 | 0.15 | -0.9223 | 0.027048 | 3 | 4.3027 | 0.067191 | 0.15 | 0.16436 | 20.0045 | 88974.1049
NG | 2025-03-13 |  | 20.0034 | 21.0386 | 20.0244 | 14.6792 | 1.1383e-05 | 88942.5926 | 19.813 | -0.95193 | 0.15
NG | 2025-03-13 |  | 20.0051 | 21.0378 | 20.0356 | 14.6781 | 1.1384e-05 | 88985.9149 | 19.8253 | -0.89893 | 0.15

## Data N2 NEL
NEL
Test Number | T03966 | [15 empty] | 0.35294
Turbine Meter
Fluid | TestPoint Timestamp | Test Point | Q_Reference | p_abs_MUT | T_MUT | rho_MUT | RE_MUT | Q_MUT | dev | U_CMC(volume) | dev average | std. dev | test points | t-distribution | uncertainty mean | U_CMC(mass) average | Total uncertainty | Q Ref average | RE MUT average
- | - | - | m3/h | bar a | °C | kg/m³ | - | m3/h | % | % | % |  |  |  | % | % | % | m3/h | -
N2 | 2024-07-09 08:31 | 1 | 397.7511 | 38.3834 | 20.1031 | 44.3514 | 3372716.8101 | 399.2938 | 0.38786 | 0.35 | 0.38224 | 0.0056957 | 3 | 4.3027 | 0.014149 | 0.35 | 0.35029 | 397.7408 | 3374042.9927
N2 | 2024-07-09 08:41 | 2 | 397.7351 | 38.382 | 20.0101 | 44.3652 | 3374385.7186 | 399.2559 | 0.38238 | 0.35
N2 | 2024-07-09 08:50 | 3 | 397.7364 | 38.3863 | 19.9971 | 44.3723 | 3375026.4495 | 399.2337 | 0.37647 | 0.35
N2 | 2024-07-09 09:02 | 4 | 332.8542 | 38.3761 | 19.9726 | 44.3646 | 2824170.1681 | 334.0196 | 0.35012 | 0.35 | 0.33153 | 0.016285 | 3 | 4.3027 | 0.040455 | 0.35 | 0.35233 | 332.918 | 2823987.2696
N2 | 2024-07-09 09:12 | 5 | 332.9493 | 38.3751 | 20.0203 | 44.3555 | 2824078.2163 | 334.014 | 0.31977 | 0.35
N2 | 2024-07-09 09:22 | 6 | 332.9504 | 38.3714 | 20.0267 | 44.3501 | 2823713.4243 | 334.0315 | 0.3247 | 0.35
N2 | 2024-07-09 09:34 | 7 | 191.2785 | 38.3609 | 20.0508 | 44.334 | 1621546.891 | 191.7823 | 0.26337 | 0.35 | 0.26116 | 0.0060128 | 3 | 4.3027 | 0.014937 | 0.35 | 0.35032 | 191.2763 | 1620957.4377
N2 | 2024-07-09 09:50 | 8 | 191.2627 | 38.3354 | 20.0194 | 44.3096 | 1620687.8401 | 191.771 | 0.26574 | 0.35
N2 | 2024-07-09 10:11 | 10 | 191.2877 | 38.3143 | 19.9596 | 44.2951 | 1620637.5821 | 191.7742 | 0.25435 | 0.35
N2 | 2024-07-09 10:26 | 11 | 95.0582 | 38.3102 | 19.9834 | 44.2864 | 805156.7406 | 95.1471 | 0.09352 | 0.35 | 0.11071 | 0.016484 | 3 | 4.3027 | 0.040949 | 0.35 | 0.35239 | 95.0351 | 804711.7342
N2 | 2024-07-09 10:46 | 12 | 95.029 | 38.298 | 19.9901 | 44.2712 | 804631.2315 | 95.1357 | 0.11223 | 0.35
N2 | 2024-07-09 11:05 | 13 | 95.0181 | 38.291 | 20.0001 | 44.2614 | 804347.2307 | 95.1382 | 0.12638 | 0.35
N2 | 2024-07-09 11:26 | 14 | 48.1913 | 38.2832 | 19.9488 | 44.2609 | 407997.77 | 48.1638 | -0.057214 | 0.35 | -0.021969 | 0.032854 | 3 | 4.3027 | 0.081615 | 0.35 | 0.35939 | 48.1739 | 407768.6534
N2 | 2024-07-09 11:45 | 15 | 48.1704 | 38.277 | 19.9417 | 44.2548 | 407775.0708 | 48.1625 | -0.016502 | 0.35
N2 | 2024-07-09 12:04 | 16 | 48.1598 | 38.2827 | 20.0261 | 44.2476 | 407533.1194 | 48.1636 | 0.0078093 | 0.35
N2 | 2024-07-09 12:25 | 17 | 24.6424 | 38.2789 | 20.0014 | 44.2472 | 208537.3771 | 24.5575 | -0.34427 | 0.35 | -0.25623 | 0.076635 | 3 | 4.3027 | 0.19037 | 0.35 | 0.39842 | 24.6103 | 208264.4249
N2 | 2024-07-09 12:45 | 18 | 24.5991 | 38.2726 | 19.9733 | 44.2446 | 208174.7136 | 24.545 | -0.21996 | 0.35
N2 | 2024-07-09 13:04 | 19 | 24.5894 | 38.2676 | 19.9615 | 44.2407 | 208081.184 | 24.5391 | -0.20447 | 0.35

## Sheet2
0 | 1 | 1.2
450 | 1 | 1.2
Graph Volume
450 | 0 | 0.35905 | NEL
470 | 0 | 0.15119 | Force
490 | 0 | 0.23396 | DNV
Graph Reynolds
7500000 | 0 | 0 | NEL
7700000 | 0 | 0 | Force
7900000 | 0 | 0 | DNV